# Activity: Test 12 Collimator Light/X-ray Field Congruence and Perpendicularity Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the diagram and image below for Collimator Light/X-ray Field Congruence and Perpendicularity Tests.

- Complete the calculations below.

- Answer the questions after the calculations.

## Evaluation A – Collimator Light/X-ray Field Congruence

- Calculate 2% of SID.

0.02 x _________cm = ________cm × 10 = _______mm

- On each side, measure and record the distance between the light field and X-ray image.

X

B

A                                     																																																					Y

A = ______mm 					X = _____mm

B = ______mm					Y = _____mm

Take the sum of the distances of opposite sides found in Step 2 and compare to the results of 2% SID found in Step 1.

A + B = _______ mm				X + Y = ______mm

A + B + X + Y = _______mm

Conclusion: Pass ______________      Fail ________________

## Acceptance Criteria:

According to SC 35, the measured misalignment of the visually defined field and the X-ray field along the length and width of the X-ray field shall not exceed 2% of the SID.

## Evaluation B – Perpendicularity

Referring to the chart below, compare and record which alignment appears most similar to the results provided.

| ~0.5 degrees | ~1.5 degrees |  |  | ~3 degrees |
| --- | --- | --- | --- | --- |
| Results: |  |  |  |  |

## Acceptance Criteria:

According to SC 35, the X-ray beam axis should be within 3 degrees of the perpendicular to the image receptor plane. Using this test tool, the upper bead should be within the outer circle to pass.

Conclusion: Pass ______________      Fail ________________

## Critical Thinking Questions

- What could cause the collimation to be off from what is set?

- Describe an alternate test that can be used to test collimator/X-ray field congruence if a collimator test template is not available.

- What could cause the beam perpendicularity to be off from 90 degrees?