Activity: Test 7 AEC Density Control Procedure

## Purpose

This activity will help you identify the procedure steps for this test.

## Instructions

- Review the images below of the equipment and materials used at Edenbrook for this test.

- Review the procedure steps from the Edenbrook MRT lab.

- Compare the textbook readings and your practicum site QA manual for similarities and differences.

## Equipments

- Dosimeter

- Two Plexiglas blocks

- Calculator

## Edenbrook MRT Procedure Steps

- Place the Plexiglass blocks (stacked) over the AEC cells of the detector ensuring the AEC cells are completely covered.

- Using a 110 cm SID, centre the dosimeter and collimate to include the Plexiglass blocks entirely.

- Place the dosimeter probe on top of the blocks to one side of the centre cell, ensuring it is not obstructing any of the AEC cells.

- Select the centre AEC cell and set 70 kVp.

- Take an exposure at each density setting: -2, -1, 0, +1 and +2 and record the mAs and mR at each setting.