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Activity: Test 19 Fluoroscopic Line Pair Resolution Procedure
Purpose
This activity will help you identify the procedure steps for this test.
Instructions
1. Review the images below of the equipment and materials used at Edenbrook for this test.
2. [bookmark: _GoBack]Review the procedure steps from the Edenbrook MRT lab.
3. Compare this procedure and your practicum site QA manual for similarities and differences between a C-arm and dedicated fluoro unit.
Equipment
· Line pair test tool
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Edenbrook MRT Procedure Steps (Note: using C-arm is limited by equipment)
1. Position the C-arm with the image intensifier on the bottom. 

2. Centre the line pair test tool on the image intensifier and collimate appropriately to include the entire test tool.

3. Using 90 kVp on the continuous fluoro setting, expose for 2 seconds and determine how far down the test tool the differences between the lines and spaces can be clearly visualized. Record the results.

The resultant image will look like this:
[image: image]
Source: Clinical Gate. Reproduced and used in accordance Fair Dealing Policy (Schedule A, Copyright of External Materials Procedure AC.2.12.1); further distribution may infringe copyright. Retrieved from http://clinicalgate.com/wp-content/uploads/2015/02/B9781437715293000373_f037-004-9781437715293.jpg

4. Repeat step 3 on both magnification settings and record the results. 

5. Change the fluoro setting to low dose and repeat steps 3 and 4. Record the results.
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