Activity: Test 3 mA Linearity Test

## Purpose:

This activity will provide you with practice in calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions:

- Review the sample data for a completed mA Linearity Test.

- Complete the calculations below.

- Answer the questions after the calculations.

- For practice purposes, you will complete only two exposures at each mA station and focal spot.

## Evaluation of Data

| Set mA | Set mAs | Focal Spot | Exposure | mR | Average mR | mR/mAs |
| --- | --- | --- | --- | --- | --- | --- |
| 50 |  | Small | 1 | 27.84 |  |  |
|  |  |  | 2 | 27.89 |  |  |
|  |  | Large | 1 | 28.24 |  |  |
|  |  |  | 2 | 28.36 |  |  |
| 100 |  | Small | 1 | 55.80 |  |  |
|  |  |  | 2 | 55.87 |  |  |
|  |  | Large | 1 | 56.77 |  |  |
|  |  |  | 2 | 56.87 |  |  |
| 125 |  | Small | 1 | 69.83 |  |  |
|  |  |  | 2 | 69.83 |  |  |
|  |  | Large | 1 | 71.02 |  |  |
|  |  |  | 2 | 71.02 |  |  |
| 160 |  | Small | 1 | 89.15 |  |  |
|  |  |  | 2 | 89.15 |  |  |
|  |  | Large | 1 | 90.57 |  |  |
|  |  |  | 2 | 90.52 |  |  |

- Complete the data table by calculating the average mR of exposures 1 and 2 for each mA setting/focal spot size. Also calculate the set mAs based on set time and mA.

- Use the average mR to calculate the mR/mAs for each of the mA/focal spot settings and record the result on the data table.

mR/mAs = Avg. mR/ set mAs

- Calculate the linearity variance of the entire mA range of the generator as follows:

|  | (Maximum mR/mAs) – (Minimum mR/mAs) | x 100 |
| --- | --- | --- |
| % Linearity = | 2 |  |
|  | Average mR/mAs |  |

= ____________________________________

- Using the “mA Linearity Worksheet” below, calculate the ratio of mR/mAs for each set of consecutive mA settings (i.e., between 50 and 100, 100 and 125, etc.) where X1 and X2 are the average mR/mAs values obtained at any two consecutive mA settings of similar focal spot size.

mA Linearity Worksheet

Small:

X1 @ 50 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 100 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

X1 @ 100 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 125 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

X1 @ 125 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 160 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

Large:

X1 @ 50 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 100 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

X1 @ 100 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 125 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

X1 @ 125 mA        (X1 – X2) ≤ 0.10(X1 + X2)

X2 @ 160 mA      (      -       ) ≤ 0.10(      +      )

_______   ≤  __________

Conclusion: Pass ______________      Fail ________________

- Compare the results to the standard in SC 35.

- The average mR/mAs values obtained at any two consecutive mA settings shall not differ more than +/- ____________ of their sum.

## Critical Thinking Questions

- What test(s) should be done prior to this one?

- What is the output of this unit being compared with?

- Will the output be the same if the kVp is changed?

- What is an indication that exposures are not in the acceptable range?