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Activity: Test 14 Lead Apparel Testing Procedure
Purpose
This activity will help you identify the procedure steps for this test.
Instructions
1. [bookmark: _GoBack]Review the images below of the equipment and materials used at Edenbrook for this test.
2. Review the procedure steps from the Edenbrook MRT lab.
3. Compare the textbook readings and your practicum site QA manual for similarities and differences.
Equipment
· Lead apparel to be tested
· Lead markers (numbers or letters)

[image: ]



Edenbrook MRT Procedure Steps 
1. Confirm that all lead apparel (aprons, gloves, thyroid collars, lead curtains, etc.) have an ID number and a permanent label stating the lead equivalency. Note if any information is missing and record.

2. Centre the tube over the table, using a 110 cm SID. Use the largest field size appropriate to cover as much of the lead apparel as possible. Start in one corner and position subsequent exposures to ensure overlap.

3. Label the images using the lead markers (numbers or letters) ensuring that the area underneath the marker has been imaged and the marker is not obscuring any crack or hole.

4. Expose apparel using 80 kVp and 5 mAs.

5. Take exposures until the entire article of lead apparel has been imaged with overlap.

6. Expose apparel using 50 kVp and 4 mAs.
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