Exam 1

Quality Assurance and Control

QUAL 370

Date:

Weight:						Marks:

Time Allotted:

Do not open this exam booklet until you are instructed to do so.

Weight:						Marks:

Time Allotted:

Do not open this exam booklet until you are instructed to do so.

Student Name:

Student ID:

Student Name:

Student ID:

# QUAL 370: Quality Assurance and Control Exam 1 Answers

### Part A: Multiple Choice Questions

What is the purpose of the AEC reproducibility test?

To ensure you have consistent kVp

To ensure you have a consistent mAs

To ensure you have consistent radiation output* Papp p. 112

To ensure you have consistent time

Which of the following tools is not used during the AEC reproducibility test?

- Dosimeter

- Lead shielding* Papp p. 112

- Acrylic blocks

- None of the above

What is the acceptable range for the AEC reproducibility test?

- All readings must be within 1% of the average

- All readings must be within 5% of the average

- All readings must be within 10% of the average

- All readings must be within 20% of the average* SC 35 p. 25

In which of the following tests do we take exposures at all of the mA stations?

- AEC density

- AEC reproducibility

- AEC sensitivity* Papp p. 111

- AEC uniformity

During the AEC density test, our mR reading should change ____% for each density setting we select.

- 5

- 10

- 15* (AHS specs)

- 20

An alternate name for the reproducibility test is:

- Coefficient of mAs

- Coefficient of kVp

- Repeatability*

- kVp accuracy

The purpose of the reproducibility test is to ensure that the unit is producing the same ______ each time.

- Exposure* Papp p. 97

- kVp

- mA

- mAs

Which of the following equipment is NOT required for the kVp accuracy test?

- kVp meter

- Dosimeter

- Densitometer*

- Lead strips

The purpose of the kVp accuracy test is to determine that the machine is producing:

- Consistent mR value for each exposure

- The kVp set on the control panel* Papp p. 98

- Consistent kVp values

- Consistent exposure times

What parameter is kept constant throughout the kVp accuracy test?

- kVp

- mAs*

- mA

- time

The relationship between mAs and dose can be described as:

- Direct*

- Indirect

- Non-threshold

- Threshold

If we double the mAs, how will the exposure be affected?

- Halved

- Doubled*

- Unable to determine

- No effect

The purpose of the timer accuracy test is to determine that the machine is:

- Exposing for the amount of time that has been set* Papp p. 101

- Producing the mAs that has been set

- Producing linear time settings

- Producing the mA that has been set

For the timer accuracy test, the results that we record are:

- Dose

- mA

- Time* Papp pp. 101–102

- kVp

As the time is halved, how is the kVp affected?

- Halved

- Doubled

- No effect*

- Quartered

Clinically, if a tube failed the perpendicularity test, a potential consequence could be:

- Missed diagnosis of an undisplaced hairline fracture* Papp p. 107

- Magnification

- Clipped anatomy

- Increased patient dose

If the coefficient of variation is smaller than the acceptable limit, what will the clinical outcome be?

- Consistent radiographic exposures attained when the same exposure factors are used* SC 35 p. 24

- Inconsistent radiographic exposures attained when the same exposure factors are use

- Consistent radiographic exposures attained when different exposure factors are used

- Inconsistent radiographic densities produced consistently

The tool used in the line pair test is called:

- Line test tool

- Line resolution tool

- Line pair test tool*

- Line pair measurement template

How does the tool used in the above mentioned test work?

- The observer counts as many lines as possible

- The observer looks for the furthest area where the lines and spaces can be differentiated*

- The observer looks for the area of fewest lines

- The observer looks for the last number seen in the space between the lines

What is the main purpose of filtration?

- To make a more homogenous beam

- To decrease patient dose by absorbing high energy x-rays

- To increase radiographic density by absorbing low energy x-rays

- To decrease patient dose by absorbing low energy x-rays* Papp p. 98

To determine the HVL reading, you:

- Divide the mean value by the original value

- Divide your original reading on the dosimeter by 2* Papp p. 99

- Calculate 20% of the original dose

- Multiply the number of aluminum sheets used by the original kVp

The purpose of the mA linearity test is to determine the relationship between _________ is linear.

- mAs and dose* Papp p. 103

- mAs and kVp

- mA and time

- mAs and filtration

If the exposure readout at 64 mAs is 400 mR, theoretically what will the exposure be at 32 mAs?

- 200 mR*

- 800 mR

- 400 mR

- 100 mR

AEC tests need to be performed:

- Monthly

- Bimonthly

- Semiannually

- Annually* SC 35 pp. 52–53

The below picture depicts the set up for which of the following tests:

| - mA linearity - Timer accuracy - AEC sensitivity - AEC field balance - Positive beam limitation - HVL |  |
| --- | --- |

All of the above

1,2,5,6

3,4,6

1,2,6*

The below picture depicts the set up for which QC tests?

- mA linearity

- Reproducibility

- AEC field balance*

- Light field-radiation field alignment

According to SC 35, the HVL at 80 kVp must be at least ____mm of aluminum.

- 2.5

- 2.9*

- 3.2

- 3.6

According to SC 35, the kVp must not deviate from the selected value by more than:

- 10%*

- 20%

- 15%

- 5%

According to SC 35, the mA must not deviate from the selected value by more than:

- 10%

- 20%*

- 15%

- 5%

According to SC 35, the loading time must not deviate from the selected value by more than:

- 10% + 1 ms*

- 20% + 1 ms

- 10% + 0.2 ms

- 20% 0.1 ms