## Slide 1: APPH 226: Post Processing and Artifacts
Objectives:
6.4, 7.1, 7.2, 7.3, 9.1, 9.2, 9.3

## Slide 2: Patient Demographic Input
Proper identification of the patient is critical
Could be impossible to retrieve images if they have not been properly identified
  Patient Name
  Facility
  Identification Numbers
  DOB
  Examination Date
  Accession Number

## Slide 3: Window and Level
Most common image processing parameters
  Adjustment of brightness and contrast
Window level
  Controls how bright/dark the image is
Window width
  Controls the contrast of the image
    Wider = lower contrast
    Narrower = higher contrast

## Slide 4: Shuttering / Masking
When viewing an image (film/digital), the unexposed edges around the image can be distractingly bright
  Causing veil glare – interferes with the image reception of the eye
With digital, automatic shuttering is used to black out the white collimation borders, eliminating veil glare
Can be manually removed
and adjusted
***This is not meant to replace collimation

## Slide 5: Image Orientation
The way that anatomy is oriented on the CR or the detector
  This is why technologists need to be particular when positioning
Image is displayed the way it is read
  CR cassettes have a label to show techs which side is up
  Sometimes upside down (ie. Table work in the single detector rooms)

## Slide 6: Image Annotation
Allows additional information to be added to the image
Film – used to write on the image with markers, place stickers
Image Annotation function with digital allows selection of preset words, or manual text input
**Should never be used as substitute for marker placement

## Slide 7: Magnification
Two types
  Magnifying glass function
    A box placed over a section of anatomy shows a magnified version of that section
    Just like using a magnifying glass to enlarge just one area
  Zoom function
    Enlarges the entire image, so that likely it will not be able to fit on the screen – only a portion will be visible at a time
    Need to use mouse to navigate the image

## Slide 8: Spatial Frequency Filtering
After signal is obtained for each pixel, to shorten the processing time and minimize storage requirements, the signals are averaged
  One pixel will be averaged with the signals of the neighboring pixels
If more pixels are involved in the averaging, the image will appear smoother
Edge enhancement is when fewer pixels are used for the averaging
  The fewer pixels are included in the average, the greater the enhancement is

## Slide 9: Spatial Frequency Filtering
High Pass Filtering
  When frequencies are amplified
  Sharpening effect
  Increases contrast and edge enhancement
Low Pass Filtering
  When frequencies are suppressed
  Smoothing
  Can lose small details, reduces noise and contrast

## Slide 10: Image Stitching
Special cases where the anatomy/area of interest is too large to fit on one cassette
Multiple images are ‘stitched’ together to create one image
  Usually 3-4 CR cassettes used
  Software program processes image to display as one single image
Eliminates need for 3-foot films – previously used in film/screen radiography
Examples: leg length, scoliosis exams

## Slide 11: Artifacts
Any undesirable density on the processed image
  Excluding scatter/fog
Four main types:
  Imaging plate artifacts
  Plate reader artifacts
  Image processing artifacts
  Printer artifacts

## Slide 12: Imaging Plate Artifacts
Usually due to aging or wear
Plate can become prone to cracks due to the constant removal and replacement within the reader
  Cracks will appear as areas of radiolucency on the image

## Slide 13: Imaging Plate Artifacts
Techs need to be careful about what goes onto the cassettes
  Adhesive tape can also leave residue on the cassette

## Slide 14: Imaging Plate Artifacts
Other Artifacts:
Static
  Causes things to stick inside the IP
  Hair/dust can cling to the plate and be visualized on image
Backscatter
  Worn down or cracked areas of the lead coating on the cassette allow scatter to reach IP
    Backscatter can be seen as dark line artifacts

## Slide 15: Image Processing Artifacts
Can occur for different reasons
  Incorrect part selection
  Incorrect sampling
Very important to set appropriate facts and choose the correct body part so that the algorithm can produce the desired image
Poor technique and positioning can cause misrepresentation of the image

## Slide 16: Plate Reader Artifacts
If there are extraneous line patterns that appear randomly (not regularly), this can be caused by issues with the electronics
  requires possible replacement of the electronics
White lines parallel to the direction of plate travel are caused by debris on the light guide
  Requires cleaning or replacement by service personnel

## Slide 17: Plate Reader Artifacts
If multiple IPs are loaded into a single cassette
  Very rare
  Only one plate will be extracted, leaving the other one to be exposed multiple times
  Similar effect as double exposing a film cassette
Insufficient erasure can result in residual information being on the IP before the next exposure
  Can occur from over exposure or if the erasure lamp requires work

## Slide 18: Plate Reader Artifacts
Moiré artifact can occur with the orientation of a stationary grid
  Grid lines need to be perpendicular to the laser scan direction
  Not an issue with moving grids

## Slide 19: Printer Artifacts
Fine white lines may appear on the image due to debris in the laser printer
This needs to be addressed by service personnel, likely just a cleaning is required

## Slide 20: Operator Error
Aside from equipment causing artifact, human error is a major contributor to artifacts on images
  Collimation
  Cassette Use
  Technical Factors

## Slide 21: Operator Errors
Collimation
  Inappropriate collimation for the body part can result in an improper calculation of the exposure indicator, affects histogram
Cassette orientation
  If the image is taken on the back of the cassette, there will be artifact from the hardware present on the back
Technical factor selection is critical to ensure image quality
  Under exposure produces quantum mottle
  Over exposure reduces contrast

## Slide 22: Operator Error
Cassette orientation
  If the image is taken on the back of the cassette, there will be artifact from the hardware present on the back
Technical factor selection is critical to ensure image quality
  Under exposure produces quantum mottle
  Over exposure reduces contrast

## Slide 23: Flat Panel Artifacts
Dead Pixels
  Manufacturing of the pixels and the array are so complex that it is conceivable that it may suffer some damage
    Can be caused by dust, scratches, static, chemicals
  Pixels can be malfunctioning or dead (no function at all)
  As detectors age, quantity of dead pixels increases, but may not need replacement if they are near edge
  Software identifies dead pixels and uses an interpolation method to ‘fill in’ the dead pixels with information from surrounding pixels
    Can be an issue if there are too many dead pixels

## Slide 24: Flat Panel Artifacts
Incorrect Gain Calibration
  Gain calibration is used for the correction of flaws in the detector
  If there is an area with many dead pixels or poor connections between the conversion layer and the array – there will be a noticeable artifact on the screen
    Gain calibration corrects this
    Creates a mask of the defects, so when an image is taken, the software uses the mask to remove the unwanted densities so only the diagnostic information is displayed

## Slide 25: Image Lag
Some flat panel systems allow the user to take images faster than the detector can accommodate
If there is an image taken before the detector has finished releasing all the signal for the previous image, there may be a faint image of the previous image still visible
Like a type of double exposure
  Caused by:
    Taking images in rapid succession
    Over exposure
    Areas with little beam attenuation (marker)

## Slide 26: Offset Correction
Offset correction can correct image lag
Determines the amount of signal inherent in the detector
  Creates a mask of the residual signal
If this is done before residual signal has left the pixels, that information can be stored as inherent and cause a negative image of the signal – inverse image

## Slide 27: Readings
Digital Radiography and PACS 2nd Edition
pg 42-51
pg 76-83
pg 91-94
Radiologic Science for Technologists 11th Ed
pg 326-329
pg 337-341
pg 355-362