APPH

226

Lab 14 – Tube Conversion – Extra Questions – ANSWER KEY

Radiographic Rating Charts and Anode Cooling Curves

Additional Questions

Complete the following questions while referring to the “Radiographic Rating Chart”

- What is the maximum kVp that can be used at 900 mA and 0.1 s? 80 kVp

- Assuming a high frequency generator, how many HU would be generated by the exposure in question 1?

HU = kVp x mA x time x modification factor

80 kVp x 900 mA x 0.1s x 1.4 = 10 080 HU

- Are the following exposure technique selections possible with this x-ray tube?

      - 120 kVp, 500 mA, 0.5 s	Yes / No

      - 80 kVp, 600 mA, 1 s	Yes / No

      - 85 kVp, 400 mA, 0.5 s	Yes / No

      - 70 kVp, 800 mA, 1s	Yes / No

      - 120 kVp, 400mA, 0.05s	Yes / No

- What is the maximum time that can be used at:

a. 110 kVp and 600 mA	d. 95 kVp and 600 mA

0.2 s	0.55 s

b. 70 kVp and 900 mA	e. 60 kVp and 500 mA

0.3 s	5.5 s

1

MRT Program

– APPH 226

c. 70 kVp and 400 mA

APPH

226

7 s

Complete the following questions while referring to the “Anode Cooling Curve”

- Assuming this is a high frequency generator, what is the maximum number of consecutive exposures that could be made, using 70 kVp and 12 mAs?

kVp x mA x time x modification factor = HU

70 kVp x 12 mAs x 1.4 = 1176 HU

  - 000 / 1176 = 297.6

297 exposures

- What is the anode cooling time needed if the anode is presently at 300 000 HU and the exposure that is needed is 120 kVp, 400 mA, 0.6 s assuming this is a three phase generator?

kVp x mA x time x modification factor = HU 120 kVp x 400 mA x 0.6s x 1.4 = 40 320 HU

  - 000 – 40 320 = 309 680 HU

2

MRT Program

– APPH 226

©

No waiting time required as the tube is presently at 300 000 HU

- What is the anode cooling time needed if the anode is presently at 325 000 HU and the exposure that is needed is 90 kVp, 800 mA, 0.8 s assuming this is a

single phase generator?

kVp x mA x time x modification factor = HU

90 kVp x 800 mA x 0.8 s x 1 = 57 600 HU

  - 000 – 57 600 =  292 400 HU

According to chart, will require ~ 20 seconds to drop from 325 000 HU to 292 400

- What is the anode cooling time needed if the anode is presently at max and the exposure that is needed is 110 kVp, 1200 mA, 0.75 s assuming this is a high frequency generator?

kVp x mA x time x modification factor = HU

110 kVp x 1200 mA x 0.75 s x 1.4 = 138 600 HU

  - 000 – 138 600 = 211 400 HU

According to chart, will require ~ 1.5 min to drop from 350 000 HU to 211 400 HU