APPH 226 – Apparatus and Image Management

# Lab 6 - Grid Errors

### Purpose

Objective: 5.3 Evaluate grids and their effects on image quality.

Recognize the cause of grid errors and understand how they can be corrected.

### Instructions

Exposure #1 – Control Image

- Enter the patient information into the console, select elbow.

- Select table detector. Set 60kVp @ 5 mAs.

- Place the grid on the table top. Ensure the “tube side” is facing up and the grid lines are running parallel to the longitudinal axis of the table.

- Set the phantom on the center of the grid.

- Center the tube to the grid, phantom & detector.

- Use an SID of 110cm and collimate to 25 x 30cm.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Exposure #2

- Maintain the same setup as Exposure #1.

- Carefully place a 15-degree sponge under the far side of the grid (aka. patients left side).

- If needed, tape the phantom in the middle of the grid, try to avoid tape on the grid.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Questions:

- What type of grid error was produced?

- How does this type of grid error affect the image quality, in comparison to the first image?

- Illustrate why the image quality is affected.

- Give an example of when this type of grid error could occur.  Explain how this error can be corrected.

Exposure #3

- Maintain the same setup as Exposure #1.

- Flip the grid over, so that the “tube side” is now facing the detector.

- Place the phantom on the center of the grid.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Questions:

- What type of grid error was produced?

- How does this type of grid error affect the image quality?

- Illustrate why the image quality is affected.

- When could this type of grid error occur in the field? Explain how this error could be corrected.

Exposure #4

- Maintain the same setup as Exposure #1.

- Place the grid so that the “tube side” is facing the tube.

- Move the grid approximately 3 inches/8 cm towards the patients left side.

- Place the phantom on the grid in the light-field.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Questions:

- What type of grid error was produced?

- How does this type of grid error affect the image quality?

- Illustrate why the image quality is affected.

- When could this type of grid error occur in the field? Explain how this error could be corrected.

- What is the term used to describe this error in positioning?

Exposure #5

- Maintain the same setup as Exposure #1.

- Rotate the grid 90 degrees so the grid lines are running transverse across the table.

- Angle the tube 15 degrees caudad.

- Center the tube to the grid and detector, using a 110cm SID.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Questions:

- What type of grid error was produced?

- Illustrate why the image quality is affected.

- When could this type of grid error occur in the field? Explain how this error could be corrected.

Exposures #6 and 7

- Maintain the same setup as Exposure #1.

- Increase SID to 140 cm.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

- Decrease the SID to 70 cm.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI/EI_s number: __________.

Questions:

- What type of grid error was produced?

- How does this type of grid error affect the image quality?

- Which exposure demonstrated greater grid cut off? Why do you think that happened?

- Illustrate why the image quality is affected.

- When could this type of grid error occur in the field? Explain how this error could be corrected (use correct terminology).

Conclusion:

- Explain which type of grid is more likely to cause an error.

- Is the orientation of the grid lines important to consider? Explain your rationale.