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Activity: Test 22 Fluoroscopic Image Lag Procedure
Purpose
This activity will help you identify the procedure steps for this test.
Instructions
1. Review the images below of the equipment and materials used at Edenbrook for this test.
2. Review the procedure steps from the Edenbrook MRT lab.
3. [bookmark: _GoBack]Compare this procedure, the textbook reading and your practicum site QA manual for similarities and differences between a C-arm and dedicated fluoro unit.
Equipment
· Coin or metal washer

[image: ]
Edenbrook MRT Procedure Steps (Note: using C-arm at Edenbrook is limited by equipment)
1. Position the C-arm with the X-ray tube on the bottom. 

2. Place a coin or washer on the table and centre it to the tube.

3. Fluoro to locate the object on the monitor.

4. While using continuous fluoro, move the C-arm around and observe the lag behind the coin or washer. It should have a comet tail-like appearance.

5. This method is subjective and acceptance is determined through the experience and impression of the observer. 
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