# Purpose

# Lab 8 - Technique Charts – ANSWER KEY

Objective: 11.3 Discuss the steps required to create a technique chart. Assess and modify technical factors.

# Instructions

Part One

- Using the APR on the console and a textbook, complete the technique chart provided.

| Part | Position | BONTRAGER |  | APR |  |
| --- | --- | --- | --- | --- | --- |
|  |  | kVp | mAs | kVp | mAs |
| Elbow | AP |  |  |  |  |
|  | Lateral |  |  |  |  |
| Shoulder | AP |  |  |  |  |
|  | Lateral (Scap Y) |  |  |  |  |
| Humerus | AP |  |  |  |  |
|  | Lateral |  |  |  |  |
| Facial Bones | Caldwell |  |  |  |  |
|  | Waters |  |  |  |  |
|  | Lateral |  |  |  |  |
|  | Townes |  |  |  |  |
| Foot | AP |  |  |  |  |
|  | Lateral |  |  |  |  |
| Wrist | PA |  |  |  |  |
|  | Oblique |  |  |  |  |
|  | Lateral |  |  |  |  |
| Knee | AP |  |  |  |  |
|  | Lateral |  |  |  |  |
| Sacrum | AP |  |  |  |  |
|  | Lateral |  |  |  |  |

# Questions:

  - Are there any differences between the techniques provided from the APR of the room compared to the techniques provided by Bontrager? If yes, why?

## For every unit/room, there will be a difference in the technical factors. Each room may be calibrated differently, and the type of generator can also affect the settings.

  - Which techniques should a technologist use?

The one for the room (using the APR) as it is designed for that room specifically.

# Instructions

Part Two

- Select a phantom, and take an image using the exposure factors from the technique chart you completed. *Remember to position the phantom appropriately to match the technique chart specifications. Use the textbook as a resource*

- Assess the image, and record the EI onto the chart below.

- Adjust the technical factors as needed. The goal is to achieve an exposure in the ideal EI range. For the purposes of this lab, aim for the exact middle of the acceptable EI range.

- Record the ideal factors into the chart below.

| Part | Position | kVp | mAs | EI | Acceptable? | New kVp | New mAs | EI |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |
|  |  |  |  |  | Yes/No |  |  |  |

- Using your knowledge about how to manipulate mAs, create a technique chart for different sized patients based on the “ideal” technical factors that you tested and recorded in step 4. Make the adjustments based on the theory regarding part thickness and technique adjustments from the powerpoint.

| Part | Position | Average Technique | Small | Large |
| --- | --- | --- | --- | --- |

For every 4-5 cm change in part thickness, the mAs should be adjusted by a factor of 2.

# Questions:

- What is the advantage of keeping the kVp constant when designing a technique chart? The advantage is that there is only one variable being changed, and mAs has a direct relationship with radiographic intensity. This makes it the easier variable to adjust because it only affects radiographic intensity and not energy. kVp will change the way that the x-rays are attenuated in the body (scatter and PE).

- How should the technique be adjusted for a hypersthenic patient?

## Increase the mAs. If the patient was very large, perhaps a small adjustment in kVp to increase penetrability of the x-ray beam, but beware of scatter.

- If a patient with osteoporosis arrived for imaging of their spine, is it advisable for the technologist to use a technique chart for their exam?

## No, in this situation, AEC would be beneficial, as we do not know the degree of the disease progression, and it is difficult to estimate what type of bone density they have.

- When could it be beneficial to utilize a technique chart instead of using AEC?

Pediatric patients – If the patient is at risk of moving, you want to set a technique so that regardless of where the part is centered, the exposure will be sufficient.

Trauma patients – Difficult to position patient appropriately with obstructed visibility (i.e. x-table c-spine in a collar). Do not take a chance repeating for poor exposure factors due to motion/poor positioning.

## *Essentially any situation in which you cannot trust that positioning will be adequate with AEC. If the part is too small, or if it is a situation where you want to minimize the chances of a repeat exposure as much as possible.

- Using information from one of the x-ray rooms in RGH, complete the following technique chart with the new technical factors to achieve optimal exposures (within the range).

Ideal Range for EI: 1800-2100

Relationship: Doubling mAs will increase EI by 300

| Part | Position | kVp | mAs | EI | Acceptable? | New kVp | New mAs | EI |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hand | PA | 52 | 2 | 1800 | Yes/No |  |  |  |
| Humerus | Lateral | 60 | 1.6 | 1700 | Yes/No | 60 | 3.2 | 2000 |
| Shoulder | Lateral | 81 | 64 | 2300 | Yes/No | 81 | 32 | 2000 |
| Foot | AP | 57 | 2.5 | 1910 | Yes/No |  |  |  |
| Knee | Obl | 66 | 3 | 1732 | Yes/No | 66 | 6 | 2032 |
| Hip | AP | 77 | 8 | 1605 | Yes/No | 77 | 16 | 1905 |
| Pelvis | AP | 77 | 10 | 1550 | Yes/No | 77 | 20 | 1850 |