Activity: Test 1 – Reproducibility/Repeatability Data Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the sample data for a completed Reproducibility/Repeatability Test.

- Complete the calculations below.

- Answer the questions after the calculations.

## Evaluation of Data

|  | Set kVp: | 80 | SID (cm): | 110 |  |
| --- | --- | --- | --- | --- | --- |
|  | Set Time (ms): | 100 | Focus: | Large |  |
|  | Set mA: | 100 |  |  |  |
|  | Exposure | mR | Time (ms) | kVp |  |
|  | 1 | 56.96 | 99.87 | 79.57 |  |
|  | 2 | 56.95 | 100.4 | 79.40 |  |
|  | 3 | 56.93 | 100.4 | 79.26 |  |
|  | 4 | 56.90 | 100.4 | 79.61 |  |
|  | 5 | 56.95 | 99.86 | 79.41 |  |
|  | 6 | 56.96 | 100.4 | 79.48 |  |
|  | 7 | 56.85 | 100.4 | 79.47 |  |
|  | 8 | 56.86 | 100.4 | 79.47 |  |
|  | 9 | 56.90 | 100.4 | 79.49 |  |
|  | 10 | 56.90 | 100.4 | 79.51 |  |
|  | Mean |  |  |  |  |
|  | % Diff Max/Min |  |  |  |  |
|  | Coefficient of Variation |  |  |  |  |

- Calculate the mean value (average) of the ten measurements and record here and on the data table.

Mean Value: mR _____ ms _____   kVp _____

- Calculate the percent difference between the mean value and the maximum value of the ten measurements as follows:

=   %

=   %

(Maximum Value – Mean Value) x 100

Mean Value

mR: ______% ms: ______%   kVp: ______%

- Calculate the percent difference between the mean value and the minimum value of the ten measurements as follows:

=    %

=    %

(Minimum Value – Mean Value) x 100

Mean Value

mR: ______% ms: ______%   kVp: ______%

- Compare the results from Step 2 to the standard in SC 35.

- Each of the ten radiation exposures measured must be within 15% of the mean value of the ten measurements.

Conclusion: Pass ______________      Fail ________________

- Calculate the coefficient of variation of the ten measurements of mR, ms and kVp using the formula below.

Reproducibility Worksheet

Step 1	Subtract each reading from the mean value, square each result and add them up.

(     -    )2 + (     -     )2 + (     -     )2 + (     -     )2 + (     -     )2

+ (     -    )2 + (     -     )2 + (     -     )2 + (     -     )2 + (     -     )2

=  _____________

(     )2 + (     )2 + (     )2 + (     )2 + (     )2 + (    )2 + (      )2 + (    )2 +

(     )2 + (     )2

____  + ____ + ____ + ____ + ____ + ____ +	 _____+ ____+ ____ + ____   = ________

Step 2	Divide the result from Step 1 by 9.

___/9 = _______

Step 3	Take the square root of the result in Step 2.

√

=  _______

Step 4	Divide the result from Step 3 by the mean value.

____ / _____ = ____

(Mean)

Step 5	Repeat for mR and kVp, then record the results on the data table.

- Compare the resultant coefficient of variation to the standard in SC 35: C < 0.05.

Conclusion: Pass ______________      Fail ________________

## Potential Difficulties

- Failure to reset the dosimeter after each exposure may result in inaccurate readings.

- All ten exposures must be obtained within a one-hour period.

- If reproducibility is greater than +/- 0.05, it could be related to either the exposure timer or mA and a service engineer would be needed to determine exact cause.

## Critical Thinking Questions

- What is the purpose of this test?

  - How might neglecting to perform this test affect the radiographer?

- What are the two formulas that are required in order to evaluate the results of this test?

- If the test were to fail, possible causes and considerations would be:

- Why would this test be difficult to perform on a falling load generator?