APPH 226 – Apparatus and Image Management

Lab 6 – AEC

Hand/Foot/Knee Phantoms, Radiation Detector and Plexiglass Blocks

### Purpose

Objective: 5.8 Discuss appropriate use of automatic exposure control (AEC).

Understand how AEC affects radiographic density.

### Instructions

Exposure #1

- Enter the patient information into the console.

- Select a technique for the phantom provided (or equivalent to utilize AEC cells, chest/abdomen).

- Set up for the table detector, using an SID of 110 cm.

- Ensure AEC is activated. Select the center cell.

- Place the dosimeter probe next to the center cell, ensuring it is not obstructing any of the AEC cell.

- Collimate to 30 x 40 cm. Place a Plexiglass block on opposite sides of the light field.

- Carefully place the phantom on top of the Plexiglass blocks, ensuring the phantom is in the middle of the light field, and there is no weight on the dosimeter.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

Dosimeter reading (exposure): _________

Exposure #2

- Maintain the same setup as Exposure #1.

- Select one side cell (right or left side).

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

Exposure: _________

Exposure #3

- Maintain the same setup as Exposure #1.

- Select one side cell and the center cell.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): __________

Exposure: __________

Discussion:

- Which specific technical factor does the AEC control?

- The patient’s T-Spine is x-rayed, and the center cell is selected. Which body part of the patient is best demonstrated when selecting AEC with the center cell for this x-ray examination?

Discussion:

- Identify the scenario(s) in which using AEC would not be appropriate.

- Imaging a wrist using the table detector with the center cell.

- An abdominal x-ray with all cells selected.

- Bilateral standing knees using the upright detector, with the center cell selected.

- T-spine x-ray using the left and center cells.

- Imaging right hip replacement, with the right cell.

- Where are AEC cells located?