## Slide 1: MRAD 251: CT Theory 1
Outcome 1 (Part 1):
Explain the basic principles of computed
tomography (CT) physics and fundamentals

## Slide 2: Objectives
Explain the process of x-ray production in an x-ray tube
Outline the basic steps to achieve a CT image
Outline the process for:
•
•
•
•
Differentiate planar imaging from tomographic imaging

## Slide 3: Objective 1.1
X.ray imaging systems are designed to control the flow of electrons, enabling a suitable beam of x-ray photons for diagnostic imaging.
Three main components of x-ray imaging systems influence the flow of electrons:
Operating Console
Generator
X-ray Tube

## Slide 4: Operating Console
A specialized keyboard, containing meters and controls for manipulating the intensity (aka. amount) and quality (aka. penetrability) of x-radiation used during imaging.
Usually located outside of the examination room
A clear protective barrier should separate the “console room” from the examination room
  Maintains patient visualization
Controls and meters are digital
  Many features are automatic

## Slide 5: Generator
Provides the power (aka. voltage) necessary to create x-ray photons in the tube at a consistent rate.
High voltage generators are preferred
  Produce a near constant-potential
    Only possess roughly 1% voltage ripple
  Produce greater x-ray quantity and quality
    Improves image quality and decreases patient dose
  Compact size

## Slide 6: X-ray Tube
An electronic vacuum that produces x-ray photons, and focuses a useful beam of radiation through a window.
Consists of 2 primary components:
  Anode
  Cathode
Is contained within a protective housing
  Helps prevent electric shock, exposure to leakage radiation, and damage to internal components
Designed to increase heat dissipation
  The production of x-rays is extremely inefficient
    99% = heat
    1% = x-ray photons

## Slide 7: X-ray Tube
The primary components (anode and cathode) are contained within a glass or metal enclosure.
Metal and Glass Envelopes:
  Maintain a vacuum inside the tube
    Allows for more efficient x-ray production and a longer tube life
    Maintains tube current
      Aka. Electron flow from cathode to anode
  Provides structural support
Metal Envelopes:
  Prevent arcing (which leads to tube failure)
    Caused by deposits from tungsten vaporization
  Enables higher tube currents
  Improves heat dissipation

## Slide 8: X-ray Tube
The window is a thin section of the metal envelope.
Emits the useful x-ray photons
Located on the Anode side
  Near the Target
Designed to enable maximum emission of x-ray photons with the least amount of absorption.

## Slide 9: Cathode
The negatively charged portion of the x-ray tube.
Purpose:
  Produce electrons
    Aka. Thermionic Emission
  Direct (aka. Focus) the electrons to a small area of the anode
Has 2 primary parts:
  Filament
  Focusing cup

## Slide 10: Cathode
Filament
Purpose:
Boil off electrons (when heated)
  When a sufficient amount of electrons have accumulated, they can be accelerated towards a metal target
Features:
X-ray tubes typically posses two filaments
  Located within close proximity to one another
  Made of thoriated tungsten
  Size of the coils used correlates to focal spot size (fss)
    Large and small fss
Only 1 filament can be used at a time
Affects spatial resolution
Impacts the heat capacity of the anode

## Slide 11: Cathode
Focusing Cup
Purpose:
Electrostatically confine electrons
  Ensures the electrons strike the anode target at the appropriate location
Features:
Made of metal
Surrounds the filament(s)
Negatively charged
  Prevents electrons from spreading out
  Directs the electrons towards the target

## Slide 12: Anode
The positively charged portion of the x-ray tube.
Purpose:
  Conduct electricity
  Produce x-rays
  Provide mechanical support for the Target
Components:
  Bearing assembly
  Rotor stud and hub
  Target (aka. Disk)
Design considerations:
  Thermal dissipation

## Slide 13: Anode
Bearing assembly
Purpose:
Provide and ensure smooth rotation of the anode disk
  Liquid bearing method improves anode disk rotation in comparison to lubricated silver ball-bearing technology
    No vibrations or noise
Conduct heat away from the x-ray tube
Rotor stud & hub
Purpose:
Prevent transmission of heat from the disk to the bearings

## Slide 14: Anode
Target
Purpose:
Collide with electrons traveling from the cathode
  X-rays are created when fast moving electrons strike a metal target
Features:
The Target is a rotating disk
  Powered by an electromagnetic induction motor
Focal Spot(s) are located on the angled portion of the target
  Area of electron collision, causing desired x-ray emission
Features are designed to improve heat capacity and spatial resolution (image
quality).

## Slide 15: CT X-ray Tube Considerations
More heat is generated during CT imaging, in comparison to general radiography.
Technology has advanced CT Tube Designs, in response to the increased heat production, and higher power levels required during imaging
 | General Radiography | Computed Tomography
 | 0.1→2.0mm | 0.5→1.0mm
 | 120→160mm | 200mm
 | 5→20o | 12o
 | ~3,400rpm | ~10,000rpm

## Slide 16: CT Targets
Basic Designs:
All-Metal disk
  Used in conventional (aka. old) scanners
  Limitations:
    Heavy
    Fixed Target
    Poor heat dissipation
    Short tube life
Brazed Graphite
  Most common design for helical/spiral CT scanners
  Tungsten-rhenium focal track & graphite base
  Advantages:
    Can withstand more heat, due to higher thermal capacity
    Longer tube life

## Slide 17: Basic Designs Continued
Chemical vapor deposition (CVD) graphite disk
  Designed for spiral/helical CT scanners
  Graphite base & very thin layer of tungsten-rhenium over the focal track
    Deposited via a chemical vapor process
  Advantages:
    Allows for large lightweight disks
    Large heat-storage capacity
    Fast cooling rates
Technological advances have increased the life of tubes significantly
From 1,000hrs→40,000hrs

## Slide 18: X-ray Production in an X-ray Tube
The primary components of an x-ray imaging system produce x- rays when:
The operating console engages the generator, sending a current of electricity to the x-ray tube
  The filament becomes heated
    Creating a source of electrons
  A potential difference develops
    Causing the electrons to accelerate towards the anode
  The target gets bombarded by electrons
    Enabling interactions which produce:
      X-ray photons
      Heat

## Slide 19: Objective 1.2
“How To” get a CT image:
Turn the equipment on, and perform QC
Prepare the examination room
Prepare the patient
Position the patient on the table
Irradiate the patient (aka. scan)
Check the images for accuracy/quality
Dismiss the patient and clean the exam room
Perform image post-processing (as needed)

## Slide 20: Basic Steps to Achieve a CT image
Computed Tomography (CT) uses large amounts of radiation. Considerations should be applied to the “basic steps”, to minimize patient dose, while maximizing diagnostic image quality.
CT Technologists must understand site protocols
  Ensures consistency when performing a CT exam
Considerations:
•
•
•
•
•
•

## Slide 21: Objective 1.3
CT is comprised of 5 main components:
Table
Gantry
Power Injector
Operating Console
Workstation
These components are used in combination to achieve a CT image in “3” steps:
Data acquisition
Image reconstruction
Image display

## Slide 22: Step 1: Data Acquisition
How we collect information to form an image
The patient is placed on a CT table
The table moves the patient in and out of the gantry aperture
  The gantry contains the x-ray tube and detectors
The x-ray tube irradiates the anatomical area of interest as the patient moves through the gantry opening
  “Scout” images are acquired prior to the CT scan
The transmitted photons are captured by the detectors
  Produces image information known as data

## Slide 23: Step 2: Image Reconstruction
How the data is manipulated prior to display
The data collected by the detectors:
Is measured
  By the DAS (data acquisition system)
    Calculates beam attenuation
    Assigns pixels different numerical values (HU)
      HU’s can be converted into different shades of gray
Needs to be converted
  Analog to Digital data conversions (ADC)
Must be processed using algorithms
  Organizes data for improved image display

## Slide 24: Step 3: Image Display
How we view Data
The reconstructed (organized) data must undergo conversion
Digital to Analog data conversion (DAC)
  Enables visualization on the Operating Console and Workstation monitor(s)
Technologists can use post-processing functions to manipulate image appearance
Image manipulations can improve image viewing
Example:
  Reformats allow the user to view anatomy in different imaging planes
  Enhances visualization of anatomical area(s) of interest

## Slide 25: Image Archive
How we store CT image data
Diagnostic quality images:
Need to be sent to a radiologist for viewing
  Aka. Image distribution via:
    PACS (picture archival communication system)
Should be accessible
  Digital images can be saved for the purpose of future viewing
    Aka. Data storage

## Slide 26: Objective 1.4
General Radiography (GR) is a form of planar imaging.
GR produces 2D images of 3D objects
  Using ionizing radiation
  Tissue attenuation differences produce image contrast
    Aka. Differential absorption
GR images display all structures contained within the part of the body irradiated.
Image limitations:
  All structures are superimposed
    Difficult to distinguish slight density changes between tissues
    Challenging to identify precise location of abnormalities

## Slide 27: Linear Tomography
The limitations of GR led to the development of tomographic imaging.
Designed to “eliminate” structures above and below an anatomical area of interest
  Structures outside the area of interest become blurred
  Improves visualization of the tissue of interest
    Irradiates a large anatomical section of the patient, however, only a small portion of the irradiated area is imaged “clearly”
    Enhances radiographic contrast
Produced by the motion of the x-ray tube and the image receptor (IR)
  Patient remains stationary
  Tube and IR motion must follow the patients longitudinal axis

## Slide 28: Linear Tomography
System design incorporates an attachment of the x-ray tube to the IR.
The tube and IR move simultaneously in opposite directions
General Concepts:
•
•
– imaginary pivot point about which the x-ray tube and the
image receptor move
No movement occurs at this point
Determines the object plane
– only section that will image anatomical structures clearly
Aka. Section Thickness or Object Plane
Size of the section imaged is affected by tube motion

## Slide 29: Linear Tomography
The tube and IR movement during exposure causes:
Objects that are positioned outside the Focal Plane to blur
  Tube movement causes these objects to exhibit varying positions on the IR
during the exposure
    An objects position on the IR correlates to tube movement
      Objects move in opposite directions
A tomographic angle
  Corresponds to the amount of tube motion during an exposure
    Decreasing tube motion, decreases tomographic angle
  Controls thickness of the section imaged

## Slide 30: Linear Tomography
Tomographic angle considerations:
Small angles create thicker sections
Angles ˂10o are defined as zonography
  Used when subject contrast is so low, large angles will not improve image quality
  Applied in chest and renal exams
0o angle would result in a section thickness of infinity
  GR applications
“Tomographic” Image Blur:
Affected by tube motion and position of anatomy
↑as distance from the Focal Plane ↑

## Slide 31: Linear Tomography
Advantages over GR:
↑Radiographic Contrast
↑Subject Contrast
↓Superimposition
Disadvantages over GR:
↑ patient dose
  X-ray tube is on for length of tube travel
  May require several exposures to ensure the tomographic section of interest is captured
  Requires the use of grids
    Only parallel grids can be used
    Grid lines must be parallel to tube movement

## Slide 32: Panoramic Tomography
Commonly used for assessment of curved bony structures.
Uses the same principles as linear tomographic imaging
Modifications:
X-ray beam is collimated to a slit
IR is slit collimated
  Using a lead shielded cover
IR translates behind the lead shielded cover
  Becomes exposed for several seconds along its length
Tube does not use oscillating motion

## Slide 33: Computed Tomography
CT was developed to further address the concerns of GR and improve the concept of linear tomography.
Minimize superimposition
Improve contrast resolution
Basic Principle:
Produce sharp, clear, cross-sectional images of the human body
  Accomplished using a rotating tube, a detector array, and computer reconstruction algorithms

## Slide 34: Computed Tomography
Applies the same principles as linear tomography.
X-ray tube and IR moves during exposure
  IR movement only applies to some CT system designs
Modifications:
The patient is moved through the system
  Images more sections of anatomy
Multiple detectors are used
  Aka. Detector array
Uses a computer to collect, and organize all the exposure data, to create an axial image display

## Slide 35: Computed Tomography
Main Advantages:
Low-contrast resolution
Versatility in data acquisition options
Image reconstruction capabilities
Reformatting possibilities
  Example: 3D imaging
Primary Disadvantages:
Radiation doses
Artifacts
↓Spatial resolution

## Slide 36: Things to take away
CT Imaging Process
  Data Acquisition
  Image Reconstruction
  Image Display
Linear Tomography
  Effects of tomographic angle
  General Concepts
    Focal Plane
    Fulcrum
  Panoramic Tomography