# APPH 226 – Apparatus and Image Management

# Worksheet - Window Width and Level

# Questions 1-10

Draw each of the gray scales & answer the following questions.

- 8 bit system    WW:  40    WL: 200

  - How many shades of gray can be displayed in this system?

  - Which pixel values will be assigned the color white?

  - Which pixel values will be assigned the color black?

  - What pixel values will be visualized at these settings?

- 10 bit system       WW:  600         WL:  700

  - How many shades of gray can be displayed in this system?

  - Which pixel values will be assigned the color white?

  - Which pixel values will be assigned the color black?

  - What pixel values will be visualized at these settings?

- 12 bit system       WW:  250         WL:  570

- How many shades of gray can be displayed in this system?

- Which pixel values will be assigned the color white?

- Which pixel values will be assigned the color black?

- What pixel values will be visualized at these settings?

- Number of Gray Levels:  4096   Area of interest has pixel values ranging from 50 to 150.

  - What is the dynamic range in bits?

  - What is the Window Width?

  - The Window Level should be set at?

  - Which pixel values will be assigned the color white?

  - Which pixel values will be assigned the color black?

- Number of Gray Levels: 256

Area of interest has pixel values ranging from 200 – 250

  - What is the dynamic range in bits?

  - What should the Window Level be set at?

  - What is the Window Width?

  - Which pixel values will be assigned the color white?

  - Which pixel values will be assigned the color black?

- Number of Gray Levels: 16 384

Area of interest has pixel values ranging from 4500-7800

- What is the dynamic range in bits?

- What should the Window Level be set at?

- What is the Window Width?

- Which pixel values will be assigned the color grey?

7.   WW:  50     WL:  500    Dynamic Range:  212

- If increased contrast is desired on the image what should be manipulated and why?

- If decreased brightness is desired on the image, what should be manipulated & what should the new setting be?

- If the dynamic range of the system was increased, would this require more computer storage space?

- How many shades of gray can be displayed with this system?

8.   WW:  370     WL:  500    Dynamic Range:  210

- If lower contrast is needed on the image what should be manipulated and why?

- If the anatomy to be visualized is lung tissue (air), what should be manipulated and why?

- How many shades of gray can be displayed with this system?

9. Two images are taken on a 12 bit system

Image 1:  WW – 100, WL - 1000

Image 2:   WW – 250 , WL – 600

  - Which image will display with higher contrast?

  - Which image will display with more brightness?

  - On Image 1, which pixel values would be assigned the color grey?

  - On Image 2, which pixel values would be assigned the color white?

10. Two images are taken on a 10 bit system

Image 1:  WW – 300, WL - 500

Image 2:   WW – 150 , WL – 200

- Which image will display with lower contrast?

- Which image will display with more brightness?

- On Image 1, which pixel values would be assigned the color black?

- On Image 2, which pixel values would be assigned the color black?