Activity: Test 13 Half Value Layer Calculations

## Purpose:

This activity will allow you to see the effect of filtration on dose/beam shaping and assess whether or not the filtration is sufficient as per SC 35.

## Instructions:

- Review the data for a completed Half Value Layer Test.

- Complete the calculations below.

- Answer the questions after the calculations.

## Evaluation of Data:

| mm Al | mR |
| --- | --- |
| 0 | 386.2 |
| 0.5 | 337.4 |
| 1 | 299.7 |
| 1.5 | 268.2 |
| 2 | 247.2 |
| 2.5 | 223.7 |
| 3 | 204.1 |
| 3.5 | 180.8 |
| 4 | 171.8 |

HVL Graph

| Exposure (mR) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 |  |
|  |  | Filtration (mm Al) |  |  |  |  |  |  |  |  |

- Plot the exposure results onto the graph and draw a “best fit” line for the data points.

- Calculate 50% of the exposure at 0 mm Al.

________ mR 0 mm Al x .05 = _________ mR (a)

- Draw a line across the graph at the level of the result (a) above.

- Determine the point where that line intersects with the best fit line and draw a line directly down from it.

- From the graph: HVL = _______ mm AL

- Compare the actual HVL to SC 35 acceptance criterion.

- According to Table 8, the SC 35 states there must be a minimum HVL of ______ mm at 80 kVp. For other X-ray tube voltages, the HVL of the radiation beam must be calculated by linear interpolation from that Table.

Conclusion: Pass ______________      Fail ________________

## Critical Thinking Questions

- What is the definition of half value layer?

- What does the HVL measure?

- What test needs to be done first in order for the HVL test to be accurate?

- Regarding inherent filtration, what is the difference between a general radiography X-ray tube and a mammographic X-ray tube?

- Name the two circumstances from which the HVL should NOT vary.