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Activity: Test 16 Fluoroscopic kVp Accuracy Procedure
Purpose
This activity will help you identify the procedure steps for this test.
[bookmark: _GoBack]Instructions
1. Review the images below of the equipment and materials used at Edenbrook for this test.
2. Review the procedure steps from the Edenbrook MRT lab.
3. Compare this procedure, the textbook reading and your practicum site QA manual for similarities and differences between a C-arm and dedicated fluoro unit.
Equipment
· Dosimeter
· Digital kVp meter (if dosimeter does not have kVp function)
· Lead apron
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Edenbrook MRT Procedure Steps (Note: using C-arm at Edenbrook is limited by equipment)
1. Position the C-arm with the image intensifier under the table. 

2. Place a lead apron or lead sheet on the table to absorb any back scatter. Position the dosimeter on top of the lead and centre it to the X-ray tube. Collimate to include the entire chamber.

3. Set the kVp to 60.

4. Expose for 2 seconds and record the reading. Repeat the exposure once more and record the reading.

5. Repeat step 4 for the following kVp settings: 70, 80, 100.
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