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Activity: Test 16 Fluoroscopy kVp Accuracy Test
Purpose
This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.
Instructions
1. [bookmark: _GoBack]Review the data for a completed Fluoro kVp Accuracy Test.
2. Complete the calculations below.
Evaluation of Data

	Set kVp
	Exposure
	kVp
	Avg kVp
	% Diff

	60
	1
	61.77
	
	

	
	2
	60.87
	
	

	70
	1
	69.74
	
	

	
	2
	69.79
	
	

	80
	1
	80.43
	
	

	
	2
	80.64
	
	

	100
	1
	102.69
	
	

	
	2
	102.92
	
	




1. Determine the percent (%) difference between set and measured kVp values. 

    % Difference = (Actual kVp) – (Set kVp) x 100
    	(Set kVp)

2. Record the % difference for each set kVp.

3. Compare the results of the % differences to the standard in SC 35.

4. The measured kVp shall have a value within +/- _____________ of the selected kVp value.

Conclusion: Pass ______________      Fail ________________
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