# Activity: Test 16 Fluoroscopy kVp Accuracy Test

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the data for a completed Fluoro kVp Accuracy Test.

- Complete the calculations below.

## Evaluation of Data

| Set kVp | Exposure | kVp | Avg kVp | % Diff |
| --- | --- | --- | --- | --- |
| 60 | 1 | 61.77 |  |  |
|  | 2 | 60.87 |  |  |
| 70 | 1 | 69.74 |  |  |
|  | 2 | 69.79 |  |  |
| 80 | 1 | 80.43 |  |  |
|  | 2 | 80.64 |  |  |
| 100 | 1 | 102.69 |  |  |
|  | 2 | 102.92 |  |  |

- Determine the percent (%) difference between set and measured kVp values.

% Difference = (Actual kVp) – (Set kVp) x 100

(Set kVp)

- Record the % difference for each set kVp.

- Compare the results of the % differences to the standard in SC 35.

- The measured kVp shall have a value within +/- _____________ of the selected kVp value.

Conclusion: Pass ______________      Fail ________________