APPH 226 – Apparatus and Image Management

Lab 5 – kVp and Exposure

Knee Phantom, Dosimeter, Lead Sheet and Plexiglass Blocks

Purpose

Objective: 5.1 Explain the prime factors and their effect on quantity and quality of the beam.

Understand how kVp affects exposure, and evaluate.

Instructions

Exposure #1

- Enter the patient information into the console, and select a knee exam.

- Select free cassette.

- Place the dosimeter on the tabletop.

- Move the table detector away to ensure it will not be exposed.

- Center the tube to the dosimeter with an SID of 110cm.

- Collimate to a 32 x 32 cm field size.

- Place one clear Plexiglass block on each side of the dosimeter (make sure the blocks are not in the light field).

- Place the phantom on top of the blocks, ensuring there is no weight on the dosimeter. (The dosimeter should be near the center of the phantom)

- Set 70kVp @ 15mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Exposure #2

- Maintain the same setup as Exposure #1.

- Find 15% of 70kVp. (10.5)

- Increase the original kVp by 15%; set: ____80___ kVp @ 15mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Exposure #3

- Maintain the same setup as Exposure #1.

- Decrease the original exposure by 15%; set:  __60____kVp, @ 15mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Discussion:

- When the kVp is increased, what happens to the quality of the photons?

- How does increasing kVp affect the quantity of transmitted photons?

Exponential Attenuation

Purpose

Objective: 3.1 Describe the effect that the human body has on attenuation.

Understand how differences in patient thickness affect exposure.

Instructions

Exposure #1

- Enter the patient information into the console, and select an abdomen exam.

- Select free cassette.

- Place the dosimeter on the tabletop.

- Move the table detector to ensure it will not be exposed.

- Center the tube to the dosimeter.

- Use a 110cm SID.

- Collimate to the outer edges of the dosimeter.

- Set 80kVp @ 25mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Exposure #2

- Maintain the same setup as Exposure #1.

- Take two glove boxes, found inside the room.  Place on each side of the dosimeter.

- Carefully place one Plexiglas block on top of the glove boxes, ensuring no weight is placed on the dosimeter.

- Set 80kVp @ 25mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Exposure #3

- Maintain the same setup as Exposure #1.

- Add another Plexiglas block on top of the first one. Ensure it is balanced, with no weight placed on the dosimeter.

- Set 80kVp @ 25mAs.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results from the dosimeter: _________ .

Discussion:

- How does increasing patient thickness affect the amount of exposure reaching the image receptor?

- How will increasing patient thickness affect the radiographic density?

- How should a technologist compensate for a larger part with mAs?  Increase, decrease or no change?