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Activity Answers: Test 16 Fluoroscopy kVp Accuracy Test 
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	Set kVp
	Exposure
	kVp
	Avg kVp
	% Diff

	60
	1
	61.77
	61.32
	2.2

	
	2
	60.87
	
	

	70
	1
	69.74
	69.77
	-0.33

	
	2
	69.79
	
	

	80
	1
	80.43
	80.535
	0.67

	
	2
	80.64
	
	

	100
	1
	102.69
	102.81
	2.81

	
	2
	102.92
	
	



1. Determine the percent (%) difference between set and measured kVp values. 

    % Difference = (Actual kVp) – (Set kVp)  x 100
    	(Set kVp)

2. Record the % difference for each set kVp.

3. Compare the results of the % differences to the standard in SC 35.

4. The measured kVp shall have a value within +/- ____10%___ of the selected kVp value.


Conclusion: Pass _____X_________      Fail ________________
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