Activity: Test 4 Timer Accuracy Test Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the sample data for a completed Timer Accuracy Test.

- Complete the calculations below.

- Answer the questions after the calculations.

## Evaluation of Data

| Set Time (ms) | Focal Spot | Result (ms) | % Difference | Pass | Fail |
| --- | --- | --- | --- | --- | --- |
| 40 | Small | 39.12 |  |  |  |
| 40 | Large | 37.63 |  |  |  |
| 50 | Small | 49.18 |  |  |  |
| 50 | Large | 47.67 |  |  |  |
| 100 | Small | 99.86 |  |  |  |
| 100 | Large | 99.39 |  |  |  |

- Complete the data table by calculating the % difference between set and measured exposure time values.

% Difference = (Actual time) - (Set time)   x 100

(Set time)

- Compare results to the standard in SC 35.

- All exposure times must be accurate to +/- _________ +1 ms of that set on the generator control panel.

## Critical Thinking Questions

- What is the purpose of this test?

- What problems might be encountered in a department if the timer has not been checked?

- When would a synchronous spinning top (such as a “Wisconsin” tool) be used? Why can the arc produced on the image be measured to give an accurate reading of the exposure time?