Guided Reading Activity: QC for X-ray Tube and Generator Answer Key

## Answers to Critical Thinking Questions

- Describe the three main parts of a quality control program for radiographic units.

  - Visual inspection

  - Environmental inspection

  - Performance testing

- What are the steps of quality control testing of a medical X-ray system?

  - Verification of system mechanical integrity and stability including safety mechanisms, auto patient release, power drives, and interlocks

  - Verification of the performance of ancillary equipment, such as film processors and display units

  - Verification of X-ray performance

  - Verification of imaging or diagnostic performance including assessment of dose

- Where should the instruction manual for the use of a radiographic system be kept?

  - Beside or very near the control panel

- Discuss the factors to evaluate within the beam restriction system and their impact on image quality and patient dose.

  - The beam restriction system is responsible for regulating the size of the X-ray field area. Because it controls the amount of patient anatomy that is exposed to radiation, it plays a significant role in patient dosage. It also plays a role in image quality and image contrast because an increase in the area of the field increases the production of scattered radiation.

- Describe the two main components of an automatic exposure control (AEC) system.

  - Detectors – also known as sensors, cells, or chambers. Actual devices can be either photodetectors, ion chambers, or solid state detectors

  - Comparator – receives signal sent by detectors and terminates exposure when a specific amount of radiation has been measured

- Discuss the various AEC system quality control parameters to be evaluated and their importance in regard to image quality.

  - Consistency of Exposure with Varying mA

    - System should maintain exposure level with any changes in mA

  - Consistency of Exposure with Varying kVp

    - System should maintain exposure with any changes in kVp

  - Consistency of Exposure with Varying Part Thickness

    - System should be able to compensate for changes in part thickness

  - Consistency of Exposure with Varying Field Sizes

    - System should be able to compensate for any changes in the X-ray field size

  - Consistency of AEC Detectors

    - Each detector or sensor should function identically to all other detectors

  - Reproducibility

    - Exposures made at the same kVp and mA stations of the same phantom thickness should produce the same exposure

  - Density Control Function

    - Density selector should allow for changes in radiation exposure of 25% to 30% for each increment

## Resources

Papp, Jeffrey. (2015). Quality management in the imaging sciences (5th ed.). St. Louis, MO: Mosby Elsevier. Retrieved from http://evolve.elsevier.com/Papp