## Slide 1: APPH 226: The Tube Objectives: 1.1-1.3

## Slide 2
“There is nothing magical about the production of x-rays; it is simply the manipulation of electricity”
Factors selected on the console to produce x-rays are:
  Volt
    Kilovolts – unit of measure for electric potential
  Ampere
    Milliampere – unit of measure for current
    Read over Chapter 4 (Bushong) – review concepts

## Slide 3: Formation of X-rays
To produce x-rays, must have:
  Source of electrons
  Means to accelerate the electrons
  Means to decelerate the electrons (target)
  Vacuum
All of these conditions are met inside of an x-ray tube

## Slide 4: Types of X-ray Tubes
Crookes tube
  Early tube – used by Roentgen
  No vacuum, gas filled
  One mA station
Coolidge tube
  Vacuum tube
  Glass or metal enclosure
  Various mA stations

## Slide 5: Tube Housing
X-ray tube is contained inside a protective housing
Must provide solid, stable support
Lead lined metal structure
  Serves as electrical insulator and thermal cushion
  Housing contains cooling oil which helps dissipate heat away from tube
Two large electrical cables enter the housing and are securely attached to the tube through special high voltage receptacles
Housing design reduces leakage radiation to acceptable limits

## Slide 6: X-Ray Tube
Relatively large in size:
  30-50 cm long
  20 cm diameter
Is an electronic vacuum tube that contains:
  _________________
  _________________
  _________________
    All contained within glass/metal envelope

## Slide 7: Glass (or metal) Envelope
Main purpose:
  Maintain a vacuum within the tube
Typically made of borosilicate glass
  Heat resistant
  Can have discoloration from vaporized tungsten coating inside on glass
    Can cause arcing and tube failure
Notes: If air is present, electrons from the air would interfere with the electron stream – cause arcing and damage tube

## Slide 8: Glass (or metal) Envelope
Metal envelope
  Maintains constant electric potential between electron stream from cathode and enclosure
  Avoids arcing problem, extends tube life
  High capacity tubes

## Slide 9: Glass (or metal) Envelope
Both enclosures have a specially designed exit point for x-rays
  Port or window
  Approximately 5 cm2 in size
Section of the enclosure that is thinner than the rest
Allows x-rays to pass through with minimal interference (absorption)

## Slide 10: Cathode Assembly
______________ side of the tube
Provides the source of electrons required for x-ray production
Made of two main parts
  Filament
  Focusing cup
Function:
  To produce or emit electrons

## Slide 11: Filament
Most tubes have two filaments
  Called dual focus tubes – large and small
Emits electrons when it is heated
Typically made of thoriated tungsten
  1-2% thorium
    Increases thermionic emission, extends filament life
  Tungsten (atomic number 74) has high melting point (3410°C),
      Doesn’t vaporize easily
  Higher thermionic emission
    Allows filament to operate at high temperatures required for x-ray production
Notes: Most common cause of tube failure is the tungsten vaporization

## Slide 12: Filament
Similar to toaster coil or lightbulb filament, but bigger
Set inside focusing cup
Diameter of filament is 1-2 mm
Length can be 1-2 cm
  Determines focal spot size
  Large or small
Typical to have two filaments parallel to each other and connected to one common ground wire

## Slide 13: Function of the Filament
Provides resistance to the flow of electrons so heat is produced
  Causes thermionic emission to occur
This happens as the exposure switch is pushed halfway down and the rotor is activated
Filament current heats up filament and electrons are boiled off by thermionic emission

## Slide 14: Focusing Cup
Made of metal (often nickel) and surrounds filament on its back and sides
  Front is open and facing anode
Negatively charged to force the electrons together into a cloud as they come off the filament – electromagnetic repulsion
Size, shape and charge of the focusing cup and how filament is situated will affect how well the electrons are ‘focused’ onto the target
Without it, electrons would repel and scatter as they come off the filament

## Slide 15: Space Charge Effect
Focusing cup shapes the electrons from filament into a cloud
  Called space charge
Self limiting
  When it reaches the size (set by current used), difficult for additional electrons to be emitted
  Saturation current – all electrons are used
This stabilizing, self limiting factor is called the ‘space charge effect’

## Slide 16: Focal Spots
Large filament = large focal spot (0.3 - 2mm)
Small filament = small focal spot (0.1 - 1 mm)
Selection made with the mA station selector on the console
  ~300 mA and under – can select either focal spot
  ~400 mA and over – only large focal spot is allowed due to heat capacity
Filament current will heat the appropriate filament as set on the console, not both

## Slide 17: Anode Assembly
____________ end of the tube
Functions:
  Electrical conductor
    Some electrons interact with target to produce x-rays, the rest continue on as current flow through the x-ray circuit
  Mechanical support for target
    Target required for electrons to interact with and produce x-rays
  Thermal dissipater
    Tremendous amount of heat generated during x-ray production

## Slide 18: Anode Design
Two designs for anodes:
_________________
  Target does not move, embedded into one end of the copper anode
  Example: dental units
  Major disadvantage – electrons always hitting same small target area, heat builds up rapidly, limited to small exposures
_________________
  This is used in general purpose rooms today
  Allows electron beam to interact with a greater target area

## Slide 19: Rotating Anode Design
Consists of:
Rotating disc made of a tungsten coating over a molybdenum core
  Alloyed Tungsten adds mechanical strength to withstand the high speed rotation and constant thermal expansion/contraction
Copper shaft with molybdenum core
  Copper is used for it’s excellent thermal and electrical conductive properties
  Molybdenum is used for the disc base and shaft core because of its low thermal conductivity and light weight strength
    Makes anode lighter and easier to rotate

## Slide 20: Rotating Anode Design
Rotation of anode increases the target area ~500 times
Heat capacity can be improved by increasing the speed of the anode rotation
  General x-ray tubes are ~3400 rpm
  High capacity x-ray tubes are ~10000 rpm

## Slide 21: Induction Motor
Anode rotates using an electromagnetic induction motor
Two major parts are:
  Stator
    _______________ the glass/metal envelope
  Rotor
    _____________the envelope
Notes: No need to engineer a motor inside the vacuum

## Slide 22: Induction Motor
Stator
  is made of electromagnetics arranged in pairs around the rotor
  OUTSIDE the enclosure
Rotor
  is a shaft made of an iron core surrounded by a copper cuff and located in the center of the stator – INSIDE the enclosure
Rotor does not touch stator, and has no electric current
  Is operated through mutual induction
Notes: Use a solid as a lubricant vs liquid because liquid would turn to gas and reduce the vacuum in tube

## Slide 23: Rotor
Inside of the rotor between the shaft and copper cuff, contains silver-plated steel ball bearings
  Helps the rotor to rotate faster
  Can hear them rotating (sound of the tube)
  Ball bearings are a high temperature lubricant
  Worn down bearings can begin to change shape, will hear a grinding sound when rotor is activated and may cause rotor to wobble
    Decreases tube efficiency

## Slide 24: Induction Motor – How it works
Stators are energized in opposing pairs, and induce a current in the rotor with an associated magnetic field
Induced field opposes that of the stator pair, and the rotor turns to correct the orientation
As the two fields align, the next pair of stators is energized, and again, another magnetic field is induced, causing the rotor to turn again

## Slide 25: Induction Motor – How it works
This continues with energizing each pair of stators in sequence so the rotor will continuously turn in to orient with the continually changing external fields
At the end of exposure, the stators will reverse direction to slow down the rotor

## Slide 26: Rotating Anode
Heat is spread out over a larger surface area
Electrons strike a small part of anode surface area, and the area changes due to the rotation
  Becomes a focal ‘track’
  Heat build up over the focal track circumference vs. one spot
Large increase of the heat load capacity and exposure techniques that can be used

## Slide 27: Rotor Mechanism Failure
If the rotor mechanism fails, this can cause the anode to overheat, resulting in:
  Pitting
  Melting
  Cracking

## Slide 28: Exposure Switch
After the switch is activated, there is a short pause before the exposure is made
During that pause while the switch is pressed:
  Rotor is accelerating
  Filament current is increasing to provide the appropriate x-ray tube current

## Slide 29: Exposure Switch
  With two position exposure switches, ideally the technologist should push the switch all the way down in one motion, avoid holding it at the first position for an extended period
    Prolongs tube life
    Minimizes unnecessary heating of the filament

## Slide 30: Putting it all together…
#1. Technologist will select the desired exposure factors on the console
  ________ – kilovolt peak
  ________– milliampere second

## Slide 31: Putting it all together….
#2. Exposure switch is pressed (half way)
  Some of the electricity goes to induction motor to bring rotor up to speed (“hitting the rotor”)
  As the rotor is coming up to speed, the selected filament is energized until the electron cloud is built up to the appropriate size
  This first phase of the exposure lasts about one second

## Slide 32: Putting it all together…
#3. Exposure switch is depressed all the way
  This is where the x-ray production is actually initialized
  Actual exposure is measured in ms
  Voltage from autotransformer (controlled by kVp selected) passes through a rectifier bank and then to the cathode and anode
    Rectifier is important so that the cathode is always the negative size and the anode is always positive
    This creates the potential difference between cathode and anode

## Slide 33: Putting it all together….
Now the things needed to produce x-rays are present…
  Source of electrons (electron cloud)
  Means to accelerate (potential difference)
  Means to decelerate the electrons (anode is ready)
  Vacuum

## Slide 34: Putting it all together…
Electron cloud is attracted to the target because of the large potential difference
  Filament electrons can reach speeds of half the speed of light in that small gap
As electrons flow from cathode to anode, this is a continuation of electricity flow through the x-ray circuit
As target is penetrated by electrons, interactions occur with the tungsten atoms
  X-rays and heat are produced

## Slide 35: Extending Tube Life
Lifespan of x-ray tube is primarily in the control of the technologist
Practices that can extend tube life:
  HEAT is the main factor/cause of tube failure
    Heat is conducted through the shaft of the anode to the rotor to the glass envelope…which then raises the temperature of the tube housing and then the surrounding air
Anode
Anode Shaft
Rotor
Glass/Metal Envelope
Oil Bath
Tube Housing

## Slide 36: Extending Tube Life
Minimum exposure factors for each exam
  Smallest exposures possible
Utilizing faster image receptors
  Minimizes exposure required
Warming up anode to prevent thermal shock
  As part of start up routine

## Slide 37: Extending Tube Life
Use shortest exposure times possible
  Extended exposure time will keep the anode at high temperature for a long time, passing heat along
  Rotor bearings can be heated excessively, increasing rotational friction and causing imbalance of the rotor anode assembly
  Continuous heat to the rotor assembly, oil bath and x-ray tube housing can cause all of these to fail
  High mA for prolonged exposure causes excessive heating of filament – more tungsten vaporized

## Slide 38: Extending Tube Life
Filament Care
  Due to high temperature of filament, the tungsten atoms are vaporized and begin to coat the inside of the enclosure
    Can lead to arcing and tube failure
    As the tungsten vaporizes, the filament wire becomes thinner and eventually breaks
      Handle tube with care

## Slide 39: Readings
Bushong: Radiologic Science For Technologists 11th Ed.
  Pg. 106-113 (end after first paragraph on right column)
  Pg 116 (start at X-Ray Tube Failure) -118 (end before Rating Charts)
  Chapter 4
    review of electricity, magnetism and electromagnetism concepts – take a quick flip through this chapter to refresh your memory