Activity: Test 10 Visual Inspection Procedure

## Purpose

This activity will help you complete a procedure checklist for a visual inspection.

## Instructions

- Complete the verification checklist for a general X-ray room in your department.  Ensure all lines are filled in to the best of your ability.

  - Consult a senior or QC technologist if there are items you encounter difficulty locating.

- Compare the textbook checklist for similarities and differences.

  - What differences do you see from the textbook compared to the Edenbrook MRT checklist?

## Equipment

- General X-ray room

## Edenbrook MRT Visual Inspection Checklist

| Compliance Verification Checklist for DI Equipment |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| This report is for : |  | Acceptance Testing |  | Compliance Testing |  |  |
| Site/Facility: |  |  |  |  |  |  |
| Room: |  |  |  |  |  |  |
| Gen. Model: |  |  |  |  |  |  |
| Manufacturer: |  |  |  |  |  |  |
| Country: |  |  |  |  |  |  |
| Serial Number: |  |  |  |  |  |  |
| Equip. Type: |  |  |  |  |  |  |
| Tube Information |  |  |  |  |  |  |
| Description: |  |  |  |  |  |  |
| Serial Number: |  |  |  |  |  |  |
| Model: |  |  |  |  |  |  |
| Date Manufactured: |  |  |  |  |  |  |
| Country Manufactured: |  |  |  |  |  |  |
| Focal Spot: | Large: ___________ mm |  | Small: _________ mm |  |  |  |
| Inherent filtration: _________ mm Al |  |  |  |  |  |  |

General Information & Facility Requirements

| Yes | No | N/A |  |
| --- | --- | --- | --- |
|  |  |  | - There is a certificate for the room. |
|  |  |  | - If “yes,” it must be posted in or nearby the room. - For mobile: if “yes,” where is it: _____________________. |
|  |  |  | - Radiation shielding is adequate for the room & meets standard. |
|  |  |  | - X-ray warning signs are posted on all entrance doors to the room. |
|  |  |  | - Protective apparel is available to patients & individuals in the radiation area. |
|  |  |  | - Does a QA program exist for the equipment? |
|  |  |  | - Does a quality maintenance program exist for the equipment? |

General Requirements for Labeling

Generator

| Yes | No | N/A |  |
| --- | --- | --- | --- |
|  |  |  | - Name of the Manufacturer |
|  |  |  | - Model Designation |
|  |  |  | - Serial Number |
|  |  |  | - Date of Manufacture |
|  |  |  | - Country of Manufacture |

X-ray Tube

| Yes | No | N/A |  |
| --- | --- | --- | --- |
|  |  |  | - Name of the Manufacturer |
|  |  |  | - Model Designation |
|  |  |  | - Serial Number |
|  |  |  | - Date of Manufacture |
|  |  |  | - Country of Manufacture |
|  |  |  | - Minimum permanent inherent filtration |
|  |  |  | - Are the anode & cathode clearly marked? |

X-ray Room

| Warning Signs |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | X-ray panel must have a permanent, visible & legible sign warning that hazardous X-rays are emitted when equipment is in operation and prohibiting unauthorized use. |
| Marking |  |  |  |  |  |
|  |  |  |  |  | All controls, lights and other indictors related to the operation of the equipment must be readily visible & clearly marked as to their function. |
| Mechanical Stability |  |  |  |  |  |
|  |  |  |  |  | - X-ray tube must be securely fixed and correctly aligned within the tube housing. |
|  |  |  |  |  | - X-ray assembly and patient support must maintain their required positions without unwarranted drift or vibration during operation. |

X-ray Room (cont’d)

| Indicators & Lights |  |  |  |
| --- | --- | --- | --- |
| - There must be distinct & separate indicators on the control panel that indicate: |  |  |  |
| Yes | No | N/A |  |
|  |  |  | - When the control panel is energized & the machine is ready to produce X-rays. |
|  |  |  | - When X-rays are being produced. |
|  |  |  | - An audible signal to indicate the end of an exposure. |
|  |  |  | - If an AEC is provided, when that mode of operation is selected. |
|  |  |  | - If the AEC mode is not selected or does not exist, the selected loading factors must be available to the operator before exposure. |
|  |  |  | - If the equipment is battery operated, indication that the battery is adequately charged. |
| - If more than one X-ray tube is controlled by one control panel: |  |  |  |
|  |  |  | A visual display on or near each X-ray tube & control panel that shows that the tube is powered. |
| Indication of Loading Factors |  |  |  |
|  |  |  | - Equipment having variable loading must have meters or other indicators on the control panel that will determine the tube voltage, current & time, or combinations of all these. |
|  |  |  | - Loading factors must be displayed before, during & after the exposure is complete. |
| Irradiation Control |  |  |  |
|  |  |  | - There must be an exposure switch, controlling timer or other mechanism to initiate & terminate the X-rays. |

| Irradiation & Emergency Switches |  |  |  |
| --- | --- | --- | --- |
| Yes | No | N/A |  |
|  |  |  | - When the equipment is equipped with an exposure switch it must require continuous pressure from the operator to produce X-rays. |
|  |  |  | - In the case of a foot switch it must prevent the production of any unintended radiation when overturned. |
|  |  |  | - In the case of mobile equipment, it must allow the operator to stand at least 3 m from the X-ray source when the tube is energized. |
|  |  |  | - If the equipment moves around the patient by remote control there must be an emergency stop switch that will stop both movement & exposure. |
| Controlling Timer |  |  |  |
| - When the equipment is equipped with a controlling timer it must be constructed so that: |  |  |  |
|  |  |  | - It can automatically terminate an exposure on completion of a preset exposure time or number of X-ray pulses. |
|  |  |  | - It permits the operator to stop the exposure at any time. |
|  |  |  | - It will automatically reset itself to its original setting or to zero at the end of the exposures. |
|  |  |  | - When it is at zero, off position or at an unmarked setting an exposure cannot occur. |
| Focal Spot Marking |  |  |  |
|  |  |  | - The focal spot must be clearly & precisely marked on the X-ray tube housing. |
|  |  |  | - Measurement from the mark to the table was taken & when compared to digital display it is accurate. |