# Activity: Test 11 Automatic Collimation Procedure

## Purpose

To confirm that the automatic collimation system (PBL or positive beam limitation) quickly and accurately adjusts to the appropriate field size of the image receptor/cassette placed in the cassette holder or bucky. This test can only be performed on equipment that has a cassette holder (bucky) where image receptors can be inserted.

## Instructions

- Review the images below of the equipment and materials used at Edenbrook for this test.

- Review the procedure steps from the Edenbrook MRT lab.

- Compare the textbook readings and your practicum site QA manual for similarities and differences.

Equipment

- Tape measure/ruler

- Calculator

## Part A:

- Using a 110 cm SID, centre the tube and cassette holder on the table. Open the light field as large as possible.

- Set the timer and time how long it takes for the light field to adjust after a cassette is inserted. Visually confirm that the light field changed to the corresponding appropriate field size of the image receptor/cassette.

- Repeat step 2 for all sizes of cassettes used in the department.

- Record the results.

## Part B:

- Using a 110 cm SID, centre the tube and cassette holder on the table.

- Place an 18 × 24 cm cassette into the cassette holder. Visually confirm the PBL.

- On the table top, centre a 24 × 30 cm cassette in the light field.

- Take an exposure at 70 kVp and 2 mAs.

- Process the 24 × 30 cm (table top) cassette and determine that the radiation field size did not exceed the image receptor by more than 3% of the SID/SDD.

- Erase the 18 × 24 cm cassette.