Final Exam

Quality Assurance and Control

QUAL 370

Date:

Final Exam

Quality Assurance and Control

QUAL 370

Date:

Weight:			30%		Marks:		ANSWERS

Time Allotted for This Exam:

Do not open this exam booklet until you are instructed to do so.

Weight:			30%		Marks:		ANSWERS

Time Allotted for This Exam:

Do not open this exam booklet until you are instructed to do so.

Student Name:

Student ID:

Student Name:

Student ID:

# QUAL 370: Quality Assurance and Control Final Exam Answers

### Multiple Choice Questions (X total marks)

One purpose for obtaining the radiation output (mR/mAs) of X-ray generators is to ensure that the:

Radiation output of three-phase and single-phase generators is similar

Reproducibility values are consistent from X-ray room to X-ray room

Percent linearity between like generators is within 5%

Radiation output of like X-ray generators is consistent from room to room* Papp p. 98

What is determined from an mA linearity test?

- A variance in mA selection will have a directly proportional effect on the exposure measured at a given kVp and time.* Papp p. 103

- Same measured exposure is obtained at a given mA selection, irrespective of mAs and kVp selected.

- A variance in mA selection will have an inversely proportional effect on the exposure measured at a given kVp and time.

- There is no known relationship between mA and measured exposure.

A radiographic room with a positive beam limiting system should not pass if the sum of the absolute values of the differences in the dimensions of the radiation field and the image receptor area exceeds ______% of the focal spot to image receptor distance.

- 2

- 3

- 4* SC 35 pp. 75–12

- None of the above

*Please use the following table to answer questions 5–10

| Table 1: Reproducibility Test Results |  |  |  |
| --- | --- | --- | --- |
| Exposure Number | Dosimeter Reading | Exposure Number | Dosimeter Reading |
| 1 | 145 mR | 6 | 148 mR |
| 2 | 144 mR | 7 | 148 mR |
| 3 | 144 mR | 8 | 145 mR |
| 4 | 149 mR | 9 | 147 mR |
| 5 | 146 mR | 10 | 150 mR |

Each of the 10 radiation exposures must be within _____% of the ______ value of the 10 measurements.

  - 10%

  - 15%

  - Mean

  - Maximum

- 1, 3

- 1, 4

- 2, 3 * SC 35 p. 24

- 2, 4

- *Please use the following table to answer questions 5–10

| Table 1: Reproducibility Test Results |  |  |  |
| --- | --- | --- | --- |
| Exposure Number | Dosimeter Reading | Exposure Number | Dosimeter Reading |
| 1 | 145 mR | 6 | 148 mR |
| 2 | 144 mR | 7 | 148 mR |
| 3 | 144 mR | 8 | 145 mR |
| 4 | 149 mR | 9 | 147 mR |
| 5 | 146 mR | 10 | 150 mR |

Each of the 10 radiation exposures must be within _____% of the ______ value of the 10 measurements.

  - 10%

  - 15%

  - Mean

  - Maximum

- 1, 3

- 1, 4

- 2, 3

- 2, 4

Part A:

- The Table 1 results are within the acceptable limits for all evaluations.

- True*

- False

Part B: Show your work below for a value of 5 marks.

Mean = 146.6 mR

Mean – max     x 100 =   146.6 – 150   x 100   =   -2.32%

Mean                                   146.6

Mean – min    x  100  =   146.6 – 144    x 100  =   1.77%

Mean                                  146.6

Both values are below 15%, which is the acceptable limit.

The purpose of calculating the coefficient of variation?

- Standard deviation

- To prove the results were not obtained by chance* SC 35 p. 24 ratio of the standard deviation to the mean value of a series of measurements

- To prove there is statistical evidence

- To prove results are within acceptable dose limits

If the coefficient of variation (C) was 0.0019, this value would be within the acceptance limit.

- True* SC 35 p. 24

- False

Results of the reproducibility test are compared to:

- Other rooms in the department

- Technique charts

- Acceptance results* Papp p. 97

- All of the above

If the coefficient of variation is greater than the acceptable limit, the outcome will be:

- Predictable radiographic exposures resulting in consistent radiographic images

- The same results attained every time the same technical factors are used

- Inconsistent radiographic exposures attained by the same technical factors*

- Different technical factors result in predictable radiographic exposures

In the clinical setting, if the central ray is NOT aligned with the centre of the detector/Bucky, a consequence could be:

- Magnification

- Clipped anatomy* Papp p. 107

- Density variation along anode-cathode axis less than 10%

- Significant anode heel effect on image

The minimum filtration required for general purpose radiographic systems at a kVp of 110 is:

- 3.2 mm Al

- 3.6 mm Al

- 3.9 mm Al* SC 35 p. 24

- 4.3 mm Al

*Please use Table 2 to answer questions 13–17

| Table 2: Leakage Radiation Test Results |  |
| --- | --- |
| Source | mR |
| Leakage | 0.22 |
| Primary | 275 |

What is the percentage of leakage radiation from the data in Table 2?

- 0.08%* activity calculation

- 1.25%

- 8.0%

- 12.5%

What is the general requirement stated in SC 35 regarding leakage radiation from X-ray tube shielding?

- Not to exceed an air kerma rate of 1.0 Gy/h

- Not to exceed an exposure rate of 115 mR/h* SC 35 p. 80 29.1

- a and b

- None of the above

When testing for leakage radiation using a paddle, measurements must be taken at a distance of 1 m away from the tabletop.

- True

- False* From the source

The entire tube housing is commonly lined with lead.

- True

- False*

Leakage radiation consists of:

- All photons escaping from the tube housing

- All photons escaping from the tube housing excluding those at the port*

- All photons escaping from the tube housing excluding those at the tabletop

- All photons escaping from the tube housing excluding those at the cathode

Which of the following tests should be performed first?

- HVL

- kVp accuracy*

- mA linearity

- Timer accuracy

An 8-penny test may also be used to evaluate:

- Collimation* Papp p. 106–107

- Perpendicularity

- Focal spot size

- Grid alignment

*Please use Table 3 to answer questions 19–22

| Table 3: mA Linearity Test Results (80 kVp and 100 ms) |  |
| --- | --- |
| mA | mR |
| 50 | 5 |
| 100 | 15 |
| 200 | 21 |
| 300 | 45 |
| 500 | 69 |
| 600 | 103 |

Calculate the percent linearity of the entire mA range of this generator.

- 1.14%

- 2.65%

- 11.3%

- 26.5%* Papp p. 103

The main purpose of this test is:

- To determine the radiation output* Papp p. 103

- To determine the amount of filtration in the tube

- To determine the accuracy of the mAs settings

- To determine the consistency of exposures

Why is this test important?

- Allows us to use linearity law

- Allows us to use repeatability law

- Allows us to use 15% rule

- Allows us to use reciprocity law* Papp p. 103

Ideally, how many exposures are required at each mA setting to obtain accurate data?

- 3

- 5

- 10*

- 15

*Please use Graph 1 to answer questions 23–27

What is the HVL for the tube tested in Graph 1?

- 2.2 mm Al

- 2.5 mm Al

- 3.2 mm Al*

- 3.5 mm Al

In accordance with the SC 35, has the minimum requirement for HVL been met for this tube?

- Yes*

- No

What is the quantity of inherent filtration in the tube?

- 2.5 mm Al equivalent

- 3.2 mm Al equivalent

- 3.5 mm Al equivalent

- There is insufficient information to determine this*

HVL is dependent on:

  - kVp

  - Total shielding

  - Type of generator

  - mAs

  - Total beam filtration

- 1,3,5*  Papp pp. 98–99

- 1,2,3

- 1,4,5

- All of the above

What values are plotted on the y-axis of the graph?

- mm Al thickness

- Exposure* Papp pp. 98–99

- mGy/mA

- Amount of filtration

Three benefits of a quality assurance program include reduced equipment down-time, decreased patient dose and decreased department morale.

- True

- False* SC 35 pp. 38–41

Who is ultimately responsible for the radiation safety of a diagnostic X-ray facility?

- The manufacturer

- The government

- The radiographer

- The owner* SC 35 p. 7

Quality assurance focuses on:

- Technical factors that can lead to variation in quality care.

- Medical factors that can lead to variations in quality care.

- Managerial factors that can lead to variations in quality care.

- Human factors that can lead to variations in quality care.* Papp p. 7

What is quality control?

- An aspect of quality assurance to optimize the performance of the operator.

- An aspect of quality assurance that monitors the technical components of the imaging systems.* Papp p. 7

- All activities that provide confidence in radiology services.

- A method of controlling the radiographic experience from start to finish.

A technologist using dosimeters in quality control testing falls into the category of:

- Level I: noninvasive and simple

- Level II: noninvasive and complex* Papp p. 7

- Level III: invasive and complex

- All of the above

Which of the following tools evaluates the relationship between the central ray and the field of exposure?

- Densitometer

- Collimator test tool

- Synchrometer

- Beam alignment test tool* Papp p. 107

According to the SC 35, cassettes must be inspected and cleaned daily.

- True

- False* SC 35 p. 45

Identify the following test pattern:

- High-contrast resolution* Papp p. 131

- Low-contrast resolution

- Line pair/mm

- Image lag

Which of the following statements regarding half value layer is true?

- HVL represents the amount of aluminum filtration in the primary beam

- HVL represents the amount of aluminum required to decrease the intensity of the primary beam by half* Papp p. 98

- A high HVL indicates that the primary beam has insufficient penetrability

- HVL must be equivalent to at least 2.5 mm of aluminum at any given kVp

The test which demonstrates that the dose remains constant between exposures when the same mA, time and kVp are selected is:

- Linearity

- Reproducibility* Papp p. 97

- Wave form accuracy

- Generator congruency

Light field–radiation field congruency can be tested:

  - Using a collimator template test tool

  - Using a nine-penny test

  - To ensure that the central ray is perpendicular to the IR

  - To ensure structures that are in the light field are included in the radiation field

- 1

- 1,2

- 1,2,4* Papp p. 105

- All of the above

Lead aprons, gloves and shields should be inspected periodically. When using fluoroscopy to test lead apparel, the cracks would appear _______ on the monitor.

- As dark areas

- As bright areas*

- Large

- As fuzzy areas

According to the SC 35, the minimum lead equivalent for protective gloves is _______ and for gonadal shielding is _________.

- 0.25 mm; 0.25 mm* SC 35 p. 35

- 0.35 mm; 0.5 mm

- 0.5 mm; 0.35 mm

- 0.25 mm; 0.5 mm

What does the linearity test determine?

- A generator must be repeatable with reasonable limits every time you change technique and then come back to the same kVp, mA and exposure time

- Measurements made under the phantom are more sensitive to changes in the kVp than those made in the air

- The output of the X-ray tube using mR/mAs for various mA and timer station combinations* Papp p. 103

- Radiographs are made with extremely low energy radiation

What is the outcome if the reproducibility results between consecutive measurements are above 20%?

- Change in radiographic density

- Change in radiographic contrast

- No change, this is within the acceptable limit

- a and b* Papp p. 97

HVL is checked routinely because:

- Softer X-rays can damage the equipment

- A decreased HVL can increase patient dose* Papp p. 98

- An increased half value layer means an increased patient dose

- The dose to the patient drops dramatically as the half value layer is decreased

What may be a consequence in the clinical setting if kVp accuracy deviates by 15%?

- There will be a change in the field congruency

- There may be a change in the image contrast* Papp p. 99

- There will be no change in patient dose

- No change, this result is within the acceptable limit

Performance testing of fluoroscopic systems must include all the following except:

- X-ray tube housing

- Automatic film processor* Papp p. 128

- Total filtration

- Automatic brightness control

Which type of test tool may be used to demonstrate pincushion distortion?

- Dosimeter

- Wire mesh tool* Papp p. 123

- Ion chamber

- Homogenous phantom

Quality control tests of fluoroscopic systems must be performed:

- Daily

- Weekly

- Semiannually

- Annually* SC 35 p. 52

At any combination of tube potential and current, fluoroscopic equipment with an automatic brightness control should not produce an air kerma rate exceeding ________ mGy/min

- 20

- 50

- 100* SC 35 p. 54 Y20

- 150

Automatic Brightness Control works by:

- Variable kVp, constant tube current

- Constant kVp, variable tube current

- Constant kVp, constant tube current

- Variable kVp, variable tube current* Papp p. 123

Automatic Gain Control works by:

- Constant kVp, variable tube current

- Constant kVp, constant tube current

- Variable kVp, variable tube current

- None of the above* Papp p. 122

Technologist responsibilities in bone mineral density include quality control through:

- Daily phantom scans*

- Generator testing

- Daily dose calibrations

- All of the above

The recommended testing frequency of kVp variance for mammographic X-ray generators is:

- Weekly

- Monthly

- Semi-annually

- Annually* SC 36 3.6 A1

The recommended testing frequency of variation for accuracy of exposure time on mammographic X-ray generators is:

- Weekly

- Monthly

- Semi-annually

- Annually* SC 36 3.6 A8

In mammography, retake analysis is performed:

- Daily

- Weekly

- Monthly*

- Semi-annually

Quality control testing on mammography equipment is performed by:

- Technologists

- QC technologist

- Medical physicist*

- All of the above

Which of the following is found inside the image receptor used for a CR system?

- Intensifying screen

- Duplitized film

- Imaging plate* Papp pp. 140–141

- All of the above

Which of the following is an advantage of CR versus film/screen radiography?

Lower patient dose

Higher-contrast resolution

Edge enhancement

All of the above*

Which of the following are components of a CR system?

  - Image receptors

  - Reader unit

  - Workstation

- 1 and 3

- 2 and 3

- 1 and 2

- 1, 2 and 3

Which of the following artifacts occurs as a result of patient motion during an exposure?

- A processing artifact

- An exposure artifact*

- A handling artifact

- No artifact

In departments with a quality management program in place, the greatest number of repeats is due to which of the following?

- Equipment problems

- Image fog

- Patient motion

- Positioning error* Papp p. 178

Which of the following artifacts can occur in both CR and film/screen radiography?

- Static electricity

- Wet-pressure sensitization

- Moire pattern* Papp p. 184

- Water spots

If a department performs 1,853 views during a 1-month period and 137 are repeated, the department repeat rate is which of the following?

- 9.37%

- 14.8%

- 7.4%* Papp p. 180

- 4.8%

Which of the following factors does NOT affect the brightness of fluoroscopic images?

- mA

- kVp

- Patient thickness and tissue density

- Time (ms)* Papp p. 123

A pattern from the Society of Motion Picture and Television Engineers can be utilized to evaluate electronic display devices.

- True* Papp p. 159

- False

Which of the following terms describes a person, department or organization that needs or wants a desired outcome?

- Supplier

- Input

- Avenger

- Customer* Papp p. 9

Examples of things covered under an operating budget for a diagnostic imaging clinic include all except:

- Technologists

- Administrative staff

- Hand sanitizer

- CR to DR conversion*

Smudge static can occur on a CR plate in very low humidity environments.

- True

- False* Papp p. 185

General radiographic X-ray tubes and mammographic X-ray tubes both use a thin glass or beryllium window which is less likely to absorb the lower energy X-rays.

- True

- False* Papp p. 195

This is an image of an AAPM TG-18 QC test pattern:

- True

- False* Papp p. 159

What could cause collimation to be out of alignment?

- Routine handling of tube housing

- Power surge through tube head

- O-ring degradation due to heat within tube housing

- Mirror shifting over time* Papp p.