APPH 226 – Apparatus and Image Management

Lab 6 – AEC

Hand and Foot Phantoms, Radiation Detector and Plexiglass Blocks

### Purpose

Objective: 5.8 Discuss appropriate use of automatic exposure control (AEC).

Understand how AEC affects radiographic density.

### Instructions

Exposure #1

- Enter the patient information into the console.

- Select a technique for the phantom provided.

- Set up for the table detector, using an SID of 110 cm.

- Ensure AEC is activated. Select the center cell.

- Place the dosimeter probe next to the center cell, ensuring it is not obstructing any of the AEC cells.

- Collimate to 30 x 40 cm. Place a Plexiglass block on opposite sides of the light field.

- Carefully place the phantom on top of the Plexiglass blocks, ensuring the phantom is in the middle of the light field, and there is no weight on the dosimeter.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

mR: _________

Exposure #2

- Maintain the same setup as Exposure #1.

- Select one side cell.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

mR: _________

Exposure #3

- Maintain the same setup as Exposure #1.

- Select the right & left cell.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

mR: _________

Exposure #4

- Maintain the same setup as Exposure #1.

- Select one side cell & the center cell.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the following readouts: mAs: _________

Time (ms): _________

mR: _________

Exposure #5

- Maintain the same setup as Exposure #1.

- Select all three cells.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the following readouts: mAs: _________

Time (ms): _________

mR: _________

Questions:

- Which technical factor(s) does the AEC control?

- Why does the use of AEC affect the dose, time and mAs?

- Which anatomy of the phantom is shown with optimum density in:

  - Exposure #1

  - Exposure #3

Exposure #6

- Maintain the same setup as Exposure #1.

- Select the center cell.

- Set the density control to -2.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the results in the table.

Exposure #7 - 10

- Repeat all the same steps as Exposure #6, but adjust the density control for each exposure as required in the data table and record results.

| Density Setting | mAs | mR |
| --- | --- | --- |
| -2 |  |  |
| -1 |  |  |
| 0 |  |  |
| +1 |  |  |
| +2 |  |  |

Questions:

- What is the approximate difference in mAs and mR between each density setting?

- When should the density control setting be utilized by the technologist?

Exposure #11

- Add a hand exam and set for table detector with an SID of 110cm.

- Select AEC with center cell only.

- Set 100 mA (except for Room 1)

- Ensure tube is centered to detector.

- Collimate to 3 x 10 cm.

- Place phantom in a lateral position, in the light field. Use a sandbag if needed.

  - Hand phantom, center cross hair should be just above shadow of thumb.

  - Foot phantom, center over metatarsals.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the following readouts

___________ mAs

___________ ms  (except for Room 1)

Exposure #12

- Maintain same setup as Exposure #11, do not adjust phantom.

- Collimate to 10 x 10 cm.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the following readouts

___________ mAs

___________ ms  (except for Room 1)

Questions:

- Was there any difference in mAs/ms between exposure 10 and 11?

- Explain the result for question 1.

- Identify the scenario(s) in which using AEC would not be appropriate.  Provide an explanation for the scenarios selected.

- Imaging a wrist using the table detector in room #2, with the center cell.

- An abdominal x-ray with all cells selected.

- Bilateral standing knees using the upright detector, with the center cell.

- T-spine x-ray using the left and center cell.

- A lateral L-spine, with a lead blocker, with the center cell.

- Imaging right hip replacement, with the right cell.

- What type of photons can be detected by the AEC cells?