## Slide 1: MRAD 251: CT Theory 1
Outcome 1 (Part 2):
Explain the basic principles of computed tomography (CT) physics and fundamentals

## Slide 2: Objective 1.3
Outline the process for:
Data Acquisition
Image Reconstruction
Image Display
Image Archive

## Slide 3: Recap: What is CT?
Computed Tomography, uses a computer to process x- ray data, taken from many different angles around a patient, to produce cross sectional images of a scanned area.
The computer quantifies all incoming x-ray data using Linear Attenuation Coefficients
  These values can be plotted on a gray scale
Enables anatomy to be visualized without superimposition

## Slide 4: Objective 1.3
There are three steps in creating a CT image:
Data Acquisition
Image Reconstruction
Image Display
  image post processing
These 3 steps describe how x-rays are processed in a CT system

## Slide 5: Data Acquisition
Refers to the method by which the patient is scanned, to achieve diagnostic images.
Scanning is defined by the beam geometry used: (size, shape, motion of the beam, and its path during the scan)
  Size: Beam is collimated, confining it to an area of interest
  Shape: Special filters shape the beam as it exits the CT x-ray tube
  Motion of the Beam: Axial or Helical
  Beam Path during scan: How it attenuated through the patient

## Slide 6: Data Acquisition
Main components that assist with the data acquisition step:
•
Aka. Donut-shaped part of the CT scanner
•
Moves the patient through the gantry
There are 2 methods of data acquisition:
Axial
Helical

## Slide 7: Axial Acquisition Method
X-ray tube rotates 3600 around the patient
  Collects attenuation data from a specified anatomical area (aka. slice)
    Used to form a single cross-sectional image
Tube stops (after the 3600 rotation)
Table moves the patient through the gantry
  Movement is a small increment
    Correlated to slice thickness
  Positions a different anatomical area under the beam
Step 1→3 is repeated until an entire body part has been scanned

## Slide 8: Axial Acquisition Method
Advantages:
Image Quality
Acquisition variability
  Contiguous
  Gapped
  Overlapped
Disadvantages:
Examination Time
  Due to inter-scan delay
Image Reconstruction capabilities

## Slide 9: Helical Acquisition Method
Aka. Spiral beam geometry, requires all three of the following to occur simultaneously:
•
slip rings
Helical scanning results in irradiating a volume of tissue.
Axial methods irradiate slice-by-slice

## Slide 10: Advantages of Helical Scanning
Reduces likelihood of misregistration
More data manipulation capabilities
Decreased scan times
  Lessens the probability of motion artifacts occurring with patients, such as:
    Uncooperative
    Trauma
    Pediatric
Enables usage of smaller volumes of contrast
  Enhances patient safety
  Improves cost effectiveness

## Slide 11: Disadvantages of Helical Scanning
Image Quality is inferior to axial acquisitions
360° of data is not obtained for each helix
  Extrapolation/Interpolation must occur
    Affects image reconstruction
* Note: Interpolation is designed to:
      Remove the “slant” from an image slice
      Create motion-free images

## Slide 12: Data Classification
All radiation that transmits through the patient and gets recorded by the detectors is known as either:
•
•
The data acquisition process produces raw data
  Which requires high storage capacities

## Slide 13: Image Reconstruction
Is the process of manipulating raw data, using algorithms.
Usually occurs automatically
  Converts the electrical signals, created by the radiation flux at the
detectors, into a digital format
Organizes raw data onto a matrix
  Using all the calculated CT Numbers
Selects the portions of data that are wanted for display

## Slide 14: Algorithms
Are a set of rules or directions used to get a specific output from a specific input.
Rules must be simple and well defined
Must terminate after a finite number of steps
Common algorithms that can be automatically applied to raw data:
  Back Projection
  Filtered Back Projection
  Iterative Reconstruction
  Fourier Transform

## Slide 15: Algorithms
Back Projection
Raw data is dragged or smeared onto a matrix, in order to view the shape of the scanned object.
Produces streak artifacts
  “Star-Pattern”
Filtered Back Projection
Uses convolution to remove the blurring caused by back projection.
Improves image quality

## Slide 16: Algorithms
Iterative Reconstruction
Compares scan data with an assumed image dataset.
Projects the calculated difference onto the matrix
Fourier Transform
Based on measuring frequencies.
Used in MRI image reconstruction

## Slide 17: More Algorithms
Many different algorithms are available for use, depending on which parts of the raw data are chosen to be enhanced or suppressed. Can be:
Selected by the technologist, or
Built into the scan protocol
  Automatically applies the algorithm(s)

## Slide 18: More Algorithms
Common algorithm selections:
Standard – Balances noise and detail
Smoothing – ↓ noise & spatial resolution, ↑ contrast resolution
  Aids visualization of soft-tissue structures, and can help minimize artifacts
•
Edge Enhancement – ↑ noise & spatial resolution, ↓ contrast resolution
  Helps demonstrate fine detail
•

## Slide 19: Image Display
The process whereby the reconstructed data is averaged, using grayscale mapping, for display on a monitor.
Using DAC (digital to analog conversion)
An image viewed on a monitor consists of data classified as .
Requires less storage capacity than raw data

## Slide 20: Image Display Considerations
Two primary categories affect how data is processed for image display:
Uncontrollable factors:
  Inherent equipment limitations
    Example: Matrix size
Controllable factors:
  Technologists selection of
    Example: Scan time

## Slide 21: Uncontrollable Factors
Matrix
A 2D array of numbers that make up a digital image
Made up of columns and rows
  Define small sections called picture elements (aka. pixels)
    Can be square or rectangular
      Shape of the pixel affects spatial resolution
The larger the matrix size, the smaller the pixel size for the same FOV
  Which improves spatial resolution

## Slide 22: Pixels
Usually have a square shape
Height of the pixel =
Width of the pixel =
axis axis
Each pixel represents a discrete value
A pixel can not display more than 1 shade at a time
  Shade derived from HU measurements
A pixels bit depth affects grayscale display
  Bit depth controls the number of different shades (from white → black) a pixel could potentially represent
    Affects contrast resolution

## Slide 23: Voxels
Used to represent a volume of tissue (aka. the scanned section thickness)
Isotropic shape is preferred
  Occurs when all three axis are equal in size
axis axis
Height of the pixel =
Width of the pixel =
Thickness of the slice = axis
Applied in 3D image displays
Each pixel is a representation of the information contained within a corresponding voxel

## Slide 24: Controllable Factors
Data acquisition can be affected by manipulating varying scan parameters, such as:
Scan field of view (SFOV)
Slice thickness
Scan time
Pitch
Other factors, such as:
Display Field of View (DFOV)
Windowing
Affect how data is processed, after data acquisition

## Slide 25: Scan Parameters
SFOV (aka. Calibration Field of View)
Determines the area within the gantry, where raw data will be acquired (aka. )
Due to the circular shape of data acquisition
Gantry is a key consideration
Identifies the number of detector cells used during data acquisition
Varying scan diameter options are available:
  Half-field or Full Field
  Small (25cm), Medium (35cm), or Large (42-50cm)
Some SFOV selections also incorporate anatomically specific processing algorithms

## Slide 26: SFOV
Scan diameters are smaller than the gantry aperture.
CT systems can not properly process any structures that are positioned outside the selected SFOV
•
Any anatomy that falls outside the selected scan diameter will result in an out-of-field artifact.

## Slide 27: Scan Parameters
Slice Thickness
Must be selected before scan acquisition.
The anatomical region of interest (ROI) should be considered when selecting the size of an acquisition slice
  Size of a scanned area will impact:
    Patient dose
    Tube heat limits
  Slice thickness affects an images:
    Spatial resolution
    Reformatting capabilities

## Slide 28: Slice Thickness Considerations
Slice thickness is also dependent on the type of system used for the scan.
SDCT (single-detector CT systems)
  Slice thickness is controlled by x-ray beam collimation
    Aka. Pre-patient collimation width
MDCT (multi-detector CT systems)
  Slice thickness is controlled by x-ray beam collimation and detector configuration

## Slide 29: Scan Parameters
Scan Time
Length of time required to irradiate an anatomical structure of interest.
Should be considered prior to scan acquisition
Operator defined parameter, that affects the outcome of the examination
  Selection of scan time may be influenced by:
    Patient condition
    Equipment limitations
      Examples: Tube heat limits; Reconstruction capabilities

## Slide 30: Scan Time Considerations
Scan time is also dependent on the type of system used, and the motion of the beam during acquisition.
Axial Scanning: irradiation is not constant
  SDCT – acquisition speed is longer, as only one slice is captured for every full rotation of the tube
  MDCT – acquisition speed is decreased, as multiple slices can be acquired for every full rotation of the tube
Helical Scanning: irradiation is constant
  Scan time is affected by table movement throughout a helical acquisition
    Speed of table movement is affected by detector configuration

## Slide 31: Scan Parameters
Pitch (P)
Describes the CT table movement throughout a helical scan acquisition.
Pitch can be defined as:
•
or
•
Pitch calculations are used to account for constant radiation output
  Calculations vary with detector configurations (MDCT vs SDCT)

## Slide 32: Pitch
Manipulations to Pitch may be required to accommodate:
Patient motion
Traumas/Emergencies
Equipment limitations
Pitch can influence anatomy coverage and scan time capabilities of an acquisition.
Pitch does not control the amount of anatomy that is required to be irradiated
  However, influences how much anatomy can be irradiated due to tube heat limits

## Slide 33: Pitch
Increasing Pitch will enable:
More anatomy (lengthwise) to be irradiated, during a shorter specified acquisition time
Shorter total scan times
Faster table movement through the gantry aperture during acquisition (aka. ↑P) will also result in:
Decreased patient motion
Improved imaging of contrast filled vasculature
Decreased patient dose

## Slide 34: Pitch
Table speed and slice thickness have a direct relationship.
When the table speed and the beam collimation are identical Pitch is expressed as:
  P = 1
This relationship can also be expressed as a ratio
  Example: a 2:1 Pitch ratio means the table will move twice the distance of the slice thickness for each rotation of the tube
    P = ?

## Slide 35: Test Your Pitch Knowledge
1. What is the pitch ratio of a 5mm section thickness with a table speed of 10mm/s?
2. What is the pitch ratio of a 20mm slice with a table speed of 10mm/s?

## Slide 36: Pitch Considerations
Pitch values selected for scan acquisition could range from 0.5 → 2.
Values greater than 1.5 will cause:
  More extrapolation
  Faster acquisition speeds
  Decreased image quality
Effects on image quality are only minimal if the Pitch does not exceed 1.5
Values less than 1 will cause:
  Scan overlap
  Slower acquisition speeds
  Increased patient dose
  Improved image quality

## Slide 37: How to calculate SDCT Pitch
SDCT Calculations:
Table movement in one gantry rotation (d)
Pitch (P) =
Slice Thickness or Beam Collimation (W)

## Slide 38: SDCT Pitch Calculation Example
Calculate pitch for a beam collimated to 10mm and a table increment of 10mm/sec.
P = d/W
P = ?
d = 10mm/sec W = 10mm
P = 1
P = 10mm/sec
10mm

## Slide 39: SDCT Pitch Practice Calculation
Calculate the Pitch of a 2.7mm section with a table feed of 4mm/s.

## Slide 40: SDCT Pitch Practice Calculation
Calculate the Pitch of a 20mm slice with a table speed of 24mm/s.

## Slide 41: MDCT Pitch Considerations
MDCT systems use a detector array
More than one detector is used to capture x-ray data
Multiple detector rows can be included within the collimated beam width
Each detector row can be used to create a slice
  Therefore, calculations are dependant on the width of detector rows

## Slide 42: How to calculate MDCT Pitch
MDCT Calculations:
Table movement in one gantry rotation (d)
Pitch (P) = Number of Slices x Slice Thickness (W)
*Note: Number of Slices x Slice Thickness = Beam width/Collimation

## Slide 43: MDCT Pitch Calculation Example
What is the Pitch of a 16-slice MDCT scanner with a 0.25mm slice thickness and table movement of 8mm per rotation?
P = d/W
P = ?
d = 8mm/sec
W = (16slices x 0.25mm)
P = 2
P = 8mm/sec
4mm

## Slide 44: MDCT Pitch Practice Calculation
Calculate the Pitch of an 8-slice MDCT system with .5mm slice thickness, and table movement of 4mm per rotation.

## Slide 45: MDCT Pitch Practice Calculation
Calculate the Pitch of a 24-slice MDCT system with .3mm slice thickness, and table movement of 11mm per rotation.

## Slide 46: Pitch and Scan Coverage
Heat is a limiting factor of scan coverage.
During helical acquisitions, the tube is emitting x-rays continually
  This produces a significant amount of heat within the CT x-ray tube
    Resulting in: Equipment Limitations
Pitch can be altered to ensure the appropriate amount of anatomy will be irradiated during scan acquisition

## Slide 47: SDCT Pitch & Scan Coverage
SDCT Calculations:
Amount of
Anatomy = P x Total acquisition time x 1/rotation time x slice thickness
covered

## Slide 48: SDCT Scan Coverage Calculations
Example:
A hypersthenic patient requires an abdominal SDCT scan.
The length of their abdomen measures 200mm from diaphragm to iliac crest
Tube heat, limits the scan time to 25 seconds
  A high mA is required due to pt. body habitus
A 2mm slice thickness is selected
Gantry speed rotation is ½ second
How much anatomy will be scanned if a 2:1 Pitch ratio is used?

## Slide 49: SDCT Scan Coverage Calculations
P = 2
Total Acquisition time = 25sec Rotation time = 0.5sec
Slice Thickness = 2mm
Amount of Anatomy Scanned = 2 x 25s x (1/0.5s) x 2mm
= 50 x 2 x 2
= 50 x 4
= 200mm
Amount of
Anatomy
covered
= P x Total acquisition time x 1/rotation time x slice thickness

## Slide 50: MDCT Pitch & Scan Coverage
MDCT Calculations:
Anatomy covered
Amount of = P x Total acquisition time x 1/rotation time x
(slice thickness x slices per rotation)

## Slide 51: MDCT Scan Coverage Calculations
Example:
A hypersthenic patient requires an abdominal CT scan, using a 4-slice multi-detector scanner.
The length of their abdomen measures 200mm from diaphragm to iliac crest
Tube heat, limits the scan time to 25 seconds
  A high mA is required due to pt. body habitus
A 1mm slice thickness is required to assess the pathology of interest
Gantry speed rotation is ½ second
How much anatomy will be scanned using a Pitch of 1?

## Slide 52: MDCT Scan Coverage Calculations
P = 1
Total Acquisition time = 25sec
Rotation time = 0.5sec Slice Thickness = 1mm
Slices per rotation = 4
Amount of Anatomy Scanned = 1 x 25s x (1/0.5s) x (1mm x 4)
x 2 x 4
= 50 x 4
= 200mm
= 25
Amount of Anatomy covered
= P x Total acquisition time x 1/rotation time x
(slice thickness x slices per rotation)

## Slide 53: Other Controllable Factors
Technologists can manipulate data, after scan acquisition.
features are applied to image data
  Example: Windowing
Exceptions:
  Some post-processing functions must be applied to raw data, retrospectively
  Examples of functions that reuse raw data:
    DFOV
    Retrospective Slice Incrementation

## Slide 54: Windowing
A post-processing technique that controls the a displayed image.
Improves visualization of a ROI
  Example: soft-tissue
&
of
Most departments have pre-programmed Window settings, due to subjectivity
Accomplished by manipulating gray scale
  Does not produce more data
  “Reassigns” a range of HU’s to each shade displayed
Comprised of two parameters:
  Window Level (WL)
  Window Width (WW)

## Slide 55: Window Level
Affects image Brightness
Determines the center value, of the range of CT numbers, in the gray scale
  Acts as a reference point
Should be centered near an average attenuation value of the tissue of interest
  To improve visualization of tissue of interest
Selects CT numbers are displayed as shades of gray
  ↑ WL = ↓ image brightness as more HU’s will display as black
  ↓ WL = ↑ image brightness as more HU’s will display as white

## Slide 56: Window Width
Affects image Contrast
Controls CT numbers are included as shades of gray for display in an image
  Represents the maximum number of gray shades that will be displayed
Any HU value that falls outside the specified WW will appear as black or white
Setting a range can improve assessment of anatomy
  ↑ WW = long gray scale = more shades of gray = ↓ contrast
  ↓ WW = short scale contrast = fewer shades of gray = ↑ contrast

## Slide 57: DFOV
Aka. Display Field of View, Zoom, or Target
Selects how much, and which section of raw data is used in image display.
Improves visualization of a region of interest (ROI)
  Influences Pixel Size – if Matrix size is unchanged, then Pixel size increases
as DFOV increases
    Affects Spatial Resolution
  A smaller DFOV produces a zoomed image
    Example: Vertebra within the SFOV of an abdominal slice
*DFOV can only be ≤ the SFOV
Localizer scans help the technologist determine and adjust the applicable DFOV size

## Slide 58: Retrospective Slice Incrementation
The ability to change the “starting point” of an image slice.
Does not affect patient dose
Only applies to Helical scans
  More manipulations can be applied to a volume of data
  Example: create overlapping slices
This post-processing feature is used because acquisition slice thickness can affect how an image is displayed.
In comparison to thin slices, thicker slices are more likely to:
  Contain superimposed structures
  Have less accurate pixel readings (as more HU’s must be averaged per pixel)
•

## Slide 59: Image Archiving
Saving images for future viewing.
Should only occur after raw data has been reconstructed, and image display is acceptable
Enables remote access to images
  Via PACS
  Using a DICOM (Digital Imaging and Communication in Medicine) standard
    Enables transfer of image data between networks

## Slide 60: Image Archiving
CT scans produce a large number of images, which are comprised of a huge amount of data.
Post-processing can decrease the amount of data required for image display
  This decreases the storage space required to “save” the image in a PACS
Accessing an electronic (softcopy) version of an image increases the variability of storage options. Examples include images:
  Stored on optical discs or DVDs
  Printed as a Hardcopy

## Slide 61: Things to take away
The difference between:
Raw Data
  Reconstruction
Image Data
  Reformatting
Post-Processing
  Retrospective Slice Incrementation
  Windowing
Pitch & Scan Coverage
  Calculations
  Consequences