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Activity: Test 21 High-Contrast (Spatial) Resolution Procedure
Purpose
This activity will help you identify the procedure steps for this test.
Instructions
1. Review the images below of the equipment and materials used at Edenbrook for this test.
2. [bookmark: _GoBack]Review the procedure steps from the Edenbrook MRT lab.
3. Compare the textbook readings and your practicum site QA manual for similarities and differences between a C-arm and dedicated fluoro unit.
Equipment
· High contrast resolution test tool
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Edenbrook MRT Procedure Steps (Note: using C-arm at Edenbrook is limited by equipment)
1. Position the C-arm with the image intensifier on the bottom. 

2. Centre the high-contrast test tool on the image intensifier and collimate appropriately to include the entire test tool.

3. Using 100 kVp on the continuous fluoro setting, expose for 2 seconds and determine the highest numbered section that the lines that can be visualized. Record the results.

The image will look like this:
[image: https://coursewareobjects.elsevier.com/objects/elr/Papp/qm5e/IC/jpg/Chapter08/008013.jpg]



4. Add one aluminum block onto the test tool and repeat the exposure. Record the results.

5. Add the second block on top of the first and expose. Record the results. 

6. Repeat steps 3–5 on both magnification settings and record the results. 

7. Change the fluoro setting to low dose, repeat steps 3–5 and record the results.
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