APPH 226 – Apparatus and Image Management

Lab 3 – The Anode Heel Effect

### Purpose

Objective: 1.5 Explain how components of the X-ray tube may affect the beam

Understand how the Anode Heel Effect relates to beam quality and demonstrate how field size and SID contribute to the Anode Heel Effect.

### Instructions

Experiment #1

- Enter patient information into the console. (Use an imaginary name, DOB, and select elbow as the body part)

- Set a technique of 65kVp @ 1.6mAs (*or 66kVp).

- Set up the room to take an exposure on the table detector with an SID of 110 cm.

- Position a body part phantom on the table in the light-field & adjust collimation to 29x43cm.

- Indicate the anode end of the image with a lead arrow & a lead “1” to identify the exposure or use available lead for identification purposes.

- Set up the dosimeter, then place the dosimeter probe in the light-field at the anode end.

- Return to the console, ensure everyone is out of the room and the door is closed.

- Notify an instructor before taking the exposure.

- Record the reading _____________.

- Move the dosimeter to the cathode end, within the light field.

- Return to the console, ensure everyone is out of the room and the door is closed

- Notify an instructor before taking the exposure.

- Record the reading ____________.

- Compare the two readings.

Conclusion:

Experiment #2

- Use the same examination entered for Experiment #1.

- Set a technique of 65kVp @ 1.6mAs (*or 66kVp).

- Set up the room to take an exposure on the table detector with an SID of 110 cm.

- Position the body part phantom on the table in the light-field & adjust collimation to 18x30cm.

- Indicate the anode end of the image with a lead arrow & a lead “2” to identify the exposure or use available lead for identification purposes.

- Place the dosimeter probe in the light-field at the anode end.

- Return to the console, ensure everyone is out of the room and the door is closed.

- Notify an instructor before taking the exposure.

- Record the reading _____________.

- Move the dosimeter to the cathode end, within the light field.

- Return to the console, ensure everyone is out of the room and the door is closed

- Notify an instructor before taking the exposure.

- Record the reading ____________.

- Compare the readings in this experiment.  Then compare with the results from Experiment #1.

Conclusion:

Experiment #3

- Use the same examination as for Experiment #1.

- Set a technique of 65kVp @ 1.6mAs (*or 66kVp).

- Set up the room to take an exposure on the table detector with an SID of 142 cm.

- Position a body part phantom on the table in the light-field & adjust collimation to 29x43cm.

- Indicate the anode end of the image with a lead arrow & a lead “3” to identify the exposure or use available lead for identification purposes.

- Place the dosimeter probe in the light-field at the anode end.

- Return to the console, ensure everyone is out of the room and the door is closed.

- Notify an instructor before taking the exposure.

- Record the reading _____________.

- Move the dosimeter to the cathode end, within the light field.

- Return to the console, ensure everyone is out of the room and the door is closed

- Notify an instructor before taking the exposure.

- Record the reading ____________.

- Compare the difference between the two readings.  Then compare with the results from Experiment #1

Conclusion:

Discussion:

- Was there a measurable difference in dose between the anode & cathode end of the 18x30cm field size?

- Was there a measurable difference in dose between the anode & cathode end of the 29x43cm field size?

- Which field size best demonstrates the anode heel effect at 110cm?

- Which SID best demonstrates the anode heel effect (110cm or 142cm) with the 29x43cm field size?

- After reviewing the images could the Anode Heel Effect have a diagnostic impact on the image. Why or Why Not?

- What is the advantage of having a designated “head of the table”?

- How could the Anode Heel Effect be used to improve the image quality of an anatomical structure?