Activity: Test 8 Automatic Exposure Control (AEC) Sensitivity Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the sample data for a completed AEC Sensitivity Test.

- Complete the calculations below.

## Evaluation of Data

| Sensitivity/Speed | mAs | mR |
| --- | --- | --- |
| S200 | 6.1 | 38.17 |
| S400 | 3.0 | 19.17 |
| S800 | 1.6 | 10.16 |

- Determine the percent of variation between speed settings, using the S400 results as the comparison point.

S200 and S400

______ % = (S200) – (S400)   x 100

(S400)

mAs        _______ % = (            ) – (           )    x 100

(         )

mR          _______  % = (           ) – (          )   x 100

(         )

S800 and S400

______ % = (S800) – (S400)   x 100

(S400)

mAs        _______ % = (            ) – (           )    x 100

(         )

mR          _______  % = (           ) – (          )   x 100

(         )

- Compare results to SC 35 and Papp, 2015, pp. 111–113.

- SC 35 states _________________________ should be used. As the speed of the system is doubled, the resultant exposure should be _______________. When the speed of the system is halved, the resultant exposure should be ______________.

Conclusion: Pass ______________      Fail ________________