Activity Answers: Test 9 Leakage Test Calculations

## Evaluation of Data

| Primary mR | 344.5 mR |
| --- | --- |
| Probe Location | Leakage mR |
| Anode side | 0.10 mR |
| Cathode Side | 0.0 mR |
| Front | 0.02 mR |
| Back | 0.0 mR |
| Above | 0.0 mR |

- Add up the individual leakage amounts to find the total Leakage mR.

(0.10 + 0.0 + 0.02 + 0.0 + 0.0) = 0.12 mR

- Determine the percent of leakage radiation using the following formula:

% = Leakage mR  x 100

Primary mR

0.12 mR    x 100		= 0.034%

344.5 mR

- Compare results to the standard in SC 35.

- SC 35 states that the leakage radiation from the X-ray source assembly shall not exceed an air kerma rate of 1.0 mGy/h at a distance of 1 m away from the focal spot used.

Conclusion: Pass ______X________      Fail ________________