Activity Answers: Test 7 Automatic Exposure Control (AEC) Density Control Calculations

## Evaluation of Data

| Density Setting | mAs | mR |
| --- | --- | --- |
| -2 | 2.0 | 12.2 |
| -1 | 2.6 | 15.57 |
| 0 | 3.2 | 19.35 |
| +1 | 4.1 | 24.14 |
| +2 | 5.0 | 30.44 |

- Determine the percent difference in mR and mAs between adjacent density control settings.

(E.g., between -2 and -1 density setting.)

% Difference = (-2 setting) – (-1 setting)  x 100

(-1 setting)

mR

(12.2) – (15.57)  x 100 = -21.64%        [Density setting -2 and -1]

(15.57)

15.57 – 19.35  x 100 = -19.53%             [Density setting -1 and 0]

19.35

19.35 – 24.14  x 100 = -19.84 %           [Density setting 0 and +1]

24.14

24.14 – 30.44  x 100 = -20.7 %

30.44				 [Density setting +1 and +2]

mAs

(2.0) – (2.6)  x 100 = -23.08 %      [Density setting -2 and -1]

(2.6)

2.6 – 3.2  x 100 = -18.75 %            [Density setting -1 and 0]

3.2

3.2 – 4.1  x 100 = -21.95%              [Density setting 0 and +1]

4.1

4.1 – 5.0  x 100 = -18 %                  [Density setting +1 and +2]

5.0

- Compare results to SC 35 and Papp, 2015, pp. 111–113.

- SC 35 states ___manufacturer’s specifications______ should be used. The AHS standard states there should be a difference of 15–25% between adjacent settings.

Conclusion: Pass _____X_________      Fail ________________