APPH

226

Worksheet - Cathode Assembly – ANSWER KEY

- What is the function of the cathode? To produce or emit electrons

- Is the cathode the positive or the negative side of the x-ray tube? Negative

- What does the cathode primarily consist of? Filament and focusing cup

- Describe the filament.

A coil of wire approximately 2 mm in diameter and 1-2 cm long.

- What happens when a filament is heated?

If heated sufficiently, thermionic emission will occur, meaning electrons will be emitted from the filament.

- The material of choice for the filament is thoriated tungsten because:

  - has a high melting point

  - higher thermionic emission

  - does not vaporize easily

- The purpose of the focusing cup is:

Confine the x-ray beam to a small area on the anode, ‘focus’ the electrons, narrow the cloud as its driven towards the anode.

- What factors can influence the effectiveness of the focusing cup?

It’s size and shape, the charge, size and shape of the filament, position of the filament within the focusing cup.

- In what situation would a dual filament unit have the filaments ‘end to end’, instead of ‘side by side’? Bi-angle x-ray tube

- When the filament current is first turned on, the tube current is also on.  True orFalse? Explain.

False, when the filament current is first turned on, it is a low current, just to warm up the filament.  No thermionic emission is occurring, so there is no tube current (no electron flow).

1

MRT

– APPH 226

©