# Activity Answers: Test 17 Maximum Fluoroscopic Exposure Rate Calculations

## Evaluation of Data

| Mode | kVp | mA | R/min | mGy/min |
| --- | --- | --- | --- | --- |
| Continuous | 110 | 3 | 1.65 | 3.1 |
| Pulse | 110 | 0.72 | 9.926 | 3.8 |
| Low Dose | 110 | 1.5 | 0.8863 | 4.6 |

Conclusion: Pass ______X________      Fail ________________