## Slide 1: APPH 226: Scatter Reduction and Beam Restriction
Objectives: 1.5, 2.3, 5.4

## Slide 2: Beam Restrictors
Device that attaches to the x-ray tube housing to regulate the size and shape of the x-ray beam.
Designed to control and minimize scatter by limiting the x-ray field size to only the anatomic structures of interest.

## Slide 3: Purpose of Beam Restriction
Benefits of restricting the beam
  Limiting patient exposure
    Large field size is not always needed
    Unnecessary skin dose
  Reducing scatter produced within patient
    Increasing contrast on the image
    Improved image quality

## Slide 4: Field Size and Collimation
Field size and collimation are synonymous
Note the different ways to describe the same thing:
  Increase collimation
  “Cone in”
  Decrease field size
And…
  Decrease collimation
  “Cone out”
  Increase field size

## Slide 5: Off-Focus Radiation
Ideally, all filament electrons will strike the anode at the focal spot
But some electrons can stray, and interact on other areas of the anode
  This creates x-rays called ‘off focus radiation’
  This type of radiation is not ideal, increases image blur, which is not ideal for image quality
    Can also image pt tissue that was meant to be excluded (i.e. ears on a skull image)
  ‘fuzzy edge’ of image – image blur/penumbra
    “zone of unsharpness”
    Like the blurry edge of a shadow
  The variable aperture collimator is designed to reduce off focus radiation

## Slide 6: Image Blur/Penumbra
Image blur/Penumbra and off-focus radiation are reduced as the beam restrictor distance is increased from the tube port.

## Slide 7: Beam Restrictors
Three types
  Aperture Diaphragm
  Cones/Cylinders
  Variable Aperture Collimator

## Slide 8: Aperture Diaphragm
Simplest
Flat piece of lead or lead lined material that fits onto the x-ray tube head
Commercially made or ‘homemade’
Opening is usually just less than the size of the IR used
Field size controlled by:
  Size of the opening
  Shape of the opening
    Can be circular, square or rectangular
  SID

## Slide 9: Aperture Diaphragm
Disadvantages
Need multiple diaphragms to achieve different field sizes
Larger area of unsharpness around image
  Due to diaphragm being located closer to the focal spot
Advantages
Inexpensive
Easy to use

## Slide 10: Cones and Cylinders
Modified aperture diaphragm
Extension cone or cylinder attached to an aperture diaphragm
Field size determined by the size and position of the distal end
Usually circular
Presently used only in special circumstances

## Slide 11: Cones
Cone has a flared extension – upper diameter smaller than lower diameter.

## Slide 12: Cylinders
Cylinders have no flaring, same diameter from top to bottom

## Slide 13: Cones and Cylinders
Disadvantages
Fixed field size
Shape of field is round
Alignment is critical between x-ray source, cone and IR
Cones-if angle of flange is greater than angle of divergence of primary beam, not effective at restricting primary beam
Advantages
  Inexpensive
  Easy to use
  Helps decrease penumbra around image (cylinder slightly more effective than cone)

## Slide 14: Image blur/Penumbra
As the distance of the beam restrictor increases from the port the more the image blur/penumbra is reduced
- cylinder is more effective at reducing image blur because it restricts the beam for a greater distance.

## Slide 15: Variable Aperture Collimator
Best and most commonly used beam restrictor
All modern general x-ray units utilize this system
Allows the most variation and flexibility in field size adjustments

## Slide 16: Variable Aperture Collimator
Two sets of shutters
  First stage - Entrance shutters
    Located right under the tube window
    Limits the x-ray beam like an aperture diaphragm
    Multiple collimator blades, on top of the collimator box
    Purpose – to control off focus radiation

## Slide 17: Variable Aperture Collimator
  Second stage shutters
    Located 3-7 inches below the tube
    Longitudinal and lateral blades
    Work in pairs that are independently controlled
      Allows for a variety of sizes
    Usually made of lead at least 3 mm thick

## Slide 18: Variable Aperture Collimator
Equipped with a light source and a mirror
Light is intended to be an accurate representation of where the primary x-ray beam will be projected during procedure
Mirror and light must be adjusted so that the projected light field coincides with the x-ray beam

## Slide 19: Variable Aperture Collimator
Collimator has a calibrated scale on the front to indicate field size
Misalignment of the mirror can cause anatomical structures to be cut off of the image
  Evaluated regularly with quality control program
  Alignment of the light field and radiation field must be within acceptable limits

## Slide 20: Alignment test – 9 penny
Any tech can perform this if they are concerned about the alignment of the light field and radiation field

## Slide 21: Positive Beam Limitation (PBL)
“Automatic Collimation”
When an IR is placed in the Bucky tray, automatic collimators mechanically adjust the size and shape of the radiation field to match that of the IR
A way of protecting patients from overexposure, but does not replace manual collimation by technologist
Purpose is to ensure the field size does not exceed that of the image receptor

## Slide 22: Positive Beam Limitation (PBL)
Still exists in digital detector system
Can be built into the APR, so when an exam is selected, the field size will automatically change
Can be over-ridden
Crucial for technologists to collimate appropriately to the area of interest, and always limited to the size of the IR

## Slide 23: Safety Code 35
States that:
“It is not sufficient to merely collimate to limit the beam to the size of the image receptor. Care should be taken to further restrict the beam to the region of the patient’s body that is of diagnostic interest.”
Safety Code 35 Section A 3.4 pg 15

## Slide 24: Compton Scatter Interaction
Scatter is produced during a Compton Interaction
Scattered photons are not a part of the useful beam
Scatter decreases the quality of the image by adding a density over the entire image that isn’t related to the anatomy
  Image fog
  No useful information

## Slide 25: Scatter and Image Quality
As scatter reaching the image increases, the image loses contrast and appears grey and dull
Only way to improve the image quality is to prevent the scatter photons from reaching the image
  There are factors within the technologist’s control that can help with this

## Slide 26: Methods of Reduction
Three primary factors that can influence the intensity of the scatter reaching the image
  kVp
  Field size
  Patient body part thickness

## Slide 27: kVp and Scatter - Review
kVp controls the energy level of the x-ray photons
When kVp is increased, the energy level is increased – this affects the interactions occurring

## Slide 28: Beam Restriction
As the x-ray beam is restricted through collimation, the amount of tissue being irradiated is decreased
  Results in:
    Improved image quality
    Reduced patient dose

## Slide 29: Beam Restriction
As the field size is decreased, less scatter radiation is produced within the patient
Less patient anatomy is being irradiated, so fewer Compton interactions are occurring
  Less scatter reaches the IR
“More matter, more scatter”

## Slide 30: Field Size and Density
When the field size is decreased, what is happening to the density on the image?
  Decreasing because…
    Fewer photons reaching IR
    Scatter, although not useful, contributes to the density of the image
    When collimating, some of the photons have been removed due to the field size – remember, the beam is not ‘squeezed’ when it is collimated

## Slide 31: Compensating for Collimation
Increase in collimation affects the number of x-rays reaching the IR to produce the latent image
Smaller field size means fewer photons hitting patient, which results in less scatter being produced
Therefore, exposure factors should be increased when collimation is increased

## Slide 32: Which exposure factor?
kVp is typically not increased
  Will increase the proportion of scatter interactions in the patient, therefore decreasing the image contrast
mAs should be altered to maintain the exposure
  Will only affect density of image

## Slide 33: Patient Thickness
The thicker the body part is, the more scatter radiation is produced
Cannot control patient thickness, but technologists can manage it by using proper technical factors and utilizing devices such as compression paddles
  Also modifications in positioning can be helpful
    PA vs AP abdomen

## Slide 34: Readings
Bushong : Radiologic Science for Technologists 11th Edition
Pg 116 (Off Focus Radiation)
Pg 187- 190 (Stop before “Control of Scatter Radiation”)
Pg 191-194 (Start at “Beam Restrictors”, and stop before “Radiographic Grids”)