RADP 215 – Radiation Protection

# Activity: Building Design and Protective Barriers

## Purpose:

This activity will enhance the student’s understanding of the factors involved in designing a building or department that will house radiology equipment.

## Instructions:

- Answer the questions below.

- Submit a completed copy into EDENBROOK.

- Your instructor will review the answers with the class.

Questions:

- List some factors that need to be taken into consideration when designing an X-ray room or department.

- What types of radiation do we need to protect against?

- Define and describe the seven factors needing consideration when determining the shielding necessary for diagnostic X-ray installations.

  - Primary or secondary barriers

  - Distance from the radiation source

  - Occupancy factor of the area (T)

| Complete the Occupancy Factors as per SC 35 |  |
| --- | --- |
| T = 1 | Administrative offices and receptionist areas, laboratories, pharmacies and other areas fully occupied by an individual, attended waiting rooms, children’s indoor play areas, adjacent X-ray control room, living quarter |
| T = 1/2 | Rooms used for patient examinations and treatments |
| T = 1/5 | Corridors, patient rooms, staff lounges, staff rest rooms |
| T = 1/8 | Corridor doors |
| T = 1/20 | Public toilets, unattended vending areas, storage rooms, outdoor areas with seating, unattended waiting rooms, patient holding areas |
| T = 1/40 | Outdoor areas with only transient pedestrian or vehicular traffic, unattended parking lots, vehicular drop off areas (unattended), attics, stairways, unattended elevators, janitor’s closets |

  - Controlled or uncontrolled areas

    - Controlled

    - Uncontrolled

  - Workload (W)

  - Maximum and average tube potential and output

  - Use factor (U)

| Complete the Use Factors as per SC 35 |  |
| --- | --- |
| Primary Barrier |  |
| U = 1 | Floors of radiation rooms, walls containing a vertical image receptor; any other walls, doors, or ceiling areas routinely exposed to the direct radiation beam |
| U = 1/4 | Doors and wall areas of radiation rooms not routinely exposed to the direct radiation beam |
| U = 1/16 | Ceiling areas of radiation rooms not routinely exposed to the direct radiation beam |
| Secondary Barrier |  |
| U = 1 | The use factor for secondary protective barriers is always taken to be 1 |