## Slide 1: RADP 215: Radiation Protection
Analyze Personnel Radiation Dose
RADP 215: Radiation Protection
Analyze Personnel Radiation Dose

## Slide 2: Radiation Units
Originally, there were three traditional measurements for radiation:
  Roentgen
  Rem
  Rad

## Slide 3: Radiation Units
Exposure in Air: a measure of the ionization of air produced by x-rays and gamma rays.
  Traditional unit = Roetgen (R)
  SI unit = C/kg
Absorbed Dose(D): measurement of energy transferred from ionizing radiation per unit mass of irradiated material
  Traditional unit = rad (rad)
  SI unit = gray (Gy)
  1 rad ~ 10mGy

## Slide 4: Radiation Units
Equivalent Dose (EqD): determined by multiplying the absorbed dose by a quality factor of the type of radiation (W(R)). Used to compare dose of different types of radiation.
EqD = D x W(R)
  X-ray has a W(R) of 1
  Traditional Unit = rem
  SI unit = Sievert (Sv)

## Slide 5: Occupational Exposure
Definition:
Traditionally expressed in rem
  (Radiation Equivalent in Man)
Sievert (Sv) has replaced rem
  1 Sv ~ 100 rem

## Slide 6: Occupational Dose
An occupational dose limit is the upper boundary dose of ionizing radiation a person is allowed to be exposed to at work.
Occupational dose limits do not include natural background radiation nor do they include radiation received for medical treatment.
millisievert (mSv) is used to measure how much radiation a person receives

## Slide 7: Dose Effects and Limits
Animal testing
Study groups
Atomic bomb survivors
  Hiroshima
  Nagasaki
Nuclear accidents
  Chernobyl
  Three Mile Island
  Japanese Tsunami

## Slide 8: Dose Effects and Limits
Radiation workers
  Early radiologists and technologists
Uranium miners
Watch dial painters (WWII)
  Radioactive paint used so dials glowed in the dark
Radiation therapy patients
  Responses to the radiation doses from treatment
Data allows results to be extrapolated for lower doses

## Slide 9: Purpose of Personnel Monitoring
Determines occupational exposure through the detection and measurement of the quantity of ionizing radiation that the dosimeter has been exposed to.
Does not protect wearer from exposure!
Awareness of occupational exposure levels

## Slide 10: Personnel Monitoring SC 35
Section 2.1.6
All operators of X-ray equipment, together with personnel (i.e., nurses) who routinely participate in radiological procedures, and others, likely to receive a radiation dose in excess of 1/20th of the dose limit to radiation workers specified in Appendix I, must be declared radiation workers and monitor their radiation exposures with the use of a personal dosimeter.

## Slide 11: Personnel Monitoring SC 35
Section 2.1.7
Personal dosimeters must be worn and stored according to the recommendations of the dosimetry service provider. When a protective apron is worn, the personal dosimeter must be worn under the apron. If extremities are likely to be exposed to significantly higher doses, additional dosimeters should be worn at those locations on the body.
Section 2.1.8
All personal dosimetry records must be maintained for the lifetime of the facility.

## Slide 12: Personnel Dosimeters
Radiation badges are lightweight and easy to carry.
Must be able to tolerate daily use.
Can detect and record both large and small exposures.
Must be reliable and consistent.
Performance cannot be influenced by outside environment.
Must be inexpensive and easy to maintain.
All must be shipped to the monitoring company for reading, so exact time of exposure cannot be determined.

## Slide 13: Film Badges
First type used
Economical and widely used in past
Consists of three parts:
  Plastic film holder
  Assortment of metal filters
  Film packet similar to dental film
Sensitive to doses from 0.1 mSv to 5,000 mSv

## Slide 14: Film Badges
Advantages:
  Film is a permanent record of exposure and can be re-read if required
  Very durable
Disadvantages:
  Sensitive to extreme temperature and humidity and can become fogged over a long period of time
Sherer Table 4-2 pg. 84 -85

## Slide 15: Thermoluminescent Dosimeter (TLD)
Appearance similar to the film badge
Light-free device containing small lithium fluoride chips functioning as sensing agents
Energized electrons are excited and get trapped in a higher level of the crystalline lattice—the “conduction band”
When read, the electrons return to a normal state, releasing visible light
Light intensity is proportional to amount of the radiation that interacted with the crystal

## Slide 16: TLD
Advantages:
  Crystals interact with X-rays in same way as human tissue
  Exposure can be measured precisely
  Environment has no affect on the reading
  Can be worn for three months and is reusable
Disadvantages:
  Higher cost
  Can only be read once and then loses information
  Recalibration required after each reading

## Slide 17: Optically Stimulated Luminescence (OSL)
Best features of the film badge and TLD
Most commonly used device
Thin-layered aluminum oxide detector
Read using a laser light
Luminesces in proportion to the radiation exposure
Can be worn for up to one year but commonly worn for quarterly readings
More sensitive than TLD (10 Sv – 10 Sv)

## Slide 18: OSL
Detector packets contain three filter blocks:
  Aluminum
  Tin
  Copper
Filters cause different degree of attenuation of radiation hitting the detector
Ideal for monitoring a worker in low radiation areas and for pregnant workers

## Slide 19: Pocket Ionization Chamber
Most sensitive type of dosimeter
Looks like fountain pen
Not commonly used in DI
Advantages:
  Immediate exposure readout for workers in high-radiation area (e.g., cardiac catheter labs)
  Exposure can be assessed after a specific procedure
  Compact, easy to carry, convenient, sensitive and reasonably accurate
Disadvantages:
  Expensive
  Must be read and zeroed everyday
  May discharge, giving false readings

## Slide 20: Other Personnel Monitors
Ring (Extremity Dosimeter)
  Worn when performing radiographic procedures that require the hands to be near the primary X-ray beam
  Interventional radiologist
Embryo-Fetus
  Additional monitor to record radiation dose to abdomen
  Provides estimate of equivalent dose to fetus

## Slide 21: Wearing Primary Dosimeters
Consistency is key
Wear it at all times
Worn on the front of the body—level depends on the area to be monitored:
  Collar level – measures dose to thyroid and neck
  Waist – measures dose to the trunk/pelvis
  Hip – measures dose to pelvis
Location and consistency is the wearer’s responsibility

## Slide 22: Primary Dosimeters
Where an apron is required:
  Fluoroscopy, interventional, OR, mobiles
Main dosimeter is worn on the front of the body, chest to waist level, facing forward
In Canada, the dosimeter is worn under the apron**
                **as per SC 35

## Slide 23: Limitations to Using Dosimeters
Accuracy of the dosimeter readings depends on safety code policies and wearer
Dosimeters must be worn whenever wearer is in a radiation area
Must be worn correctly

## Slide 24: Radiation Safety for Staff
Time – in room only when necessary
Distance – Inverse Square Law
Staff shielding –
  Aprons
  Gloves
  Thyroid collars
  Lead glasses
  Lead doors
Structural barriers – walls, windows

## Slide 25: Radiation Dosimetry Reports
Health Canada’s National Dosimetry Services (NDS) and Radiation Protection Bureau oversees radiation exposure records.
Maintain a permanent record of radiation exposure for primary radiation workers.
Reports are issued quarterly by Dosimetry Services and are posted in the DI department.
Every radiation worker is on the report and individuals are expected to review the report for their reading.
If a readout is registered, a change in work habits should be investigated.

## Slide 26: National Dose Registry (NDR)
The National Dose Registry (NDR) contains the dose records of individuals who are monitored for occupational exposures to ionizing radiation.
The National Dose Registry will release a Personal Dose History Summary (PDHS) upon request. Employers and prospective employers can obtain an Employee's Dose History Summary (EDHS) upon written consent of the radiation worker.