## Slide 1: APPH 226: Digital Image Acquisition : PSP (Cassette based) and the Reader
Objectives: 4.1, 4.2, 4.3, 6.1

## Slide 2: Digital Radiography
Categorized into:
Computed radiography
  CR
  PSP (photostimulable phosphors)
Direct radiography
  DR
  FPD (flat panel detectors)

## Slide 3: Computed Radiography
Can be integrated into existing radiography rooms
No physical modification of x-ray unit required
  AEC may need to be re-calibrated
  Techniques within APR and manual technique charts will need to be adjusted
Require a CR reader and a computer workstation

## Slide 4: Computed Radiography
Main difference between PSP and conventional radiography is:
  An imaging plate (IP) takes the place of film and is in charge of storing the image
  After the exposure, the cassette is ‘read’ and the phosphors release the stored light, converting it into an electrical signal, which is then digitized
  The cassette is simply a lightproof container which protects an imaging plate instead of film

## Slide 5: Primary Parts of CR system
Cassette
PSP plate (inside cassette)
Cassette/CR reader
Computer work station

## Slide 6: Cassette
Purpose: to contain the PSP plate
Made of lightweight plastic
Inside is lined with felt to protect against static buildup, dust and mechanical damage to the plate
Backing made of aluminum or lead to absorb backscatter radiation
Will have a window for barcode label to scan cassette to match with image information

## Slide 7: PSP Plate or Imaging Plate (IP)
Made up of several layers, similar to film-screen based system
  Protective layer
  Phosphor layer
  Reflective layer
  Conductive layer
  Support layer

## Slide 8: Layers of Imaging Plate
Protective Layer
  Purpose: to protect the phosphor layer
  Thin layer of plastic
  Coating over top of the phosphor layer

## Slide 9: Layers of PSP plate
Phosphor Layer (Active Layer)
  Critical component of plate, contains the phosphors
  Usually barium fluorohalide with europium
  May contain dye to help prevent light spread
  “Traps” electrons during exposure
  Can be turbid or structured
    Turbid – random distribution of phosphor crystals within active layer
    Structured – columnar phosphor crystals within active layer resembling needles standing on end and packed together

## Slide 10: Layers of PSP Plate
Reflective Layer
  Reflects light released during ‘reading phase’ toward the photodetector
Conductive Layer
  Reduces and conducts away static electricity
Support Layer
  Sturdy material to give some structure to the plate
  May also have lead layer to absorb backscatter

## Slide 11: Barcode
On backing of cassette, there will be a window with a barcode
Allows technologist to match the cassette with the patient information
For every image, the barcode must be scanned to ‘assign’ the correct patient information to it

## Slide 12: Acquiring the Image
Image acquisition occurs in the same way
A patient is positioned, and the body part aligned with the IR
Patient is exposed with proper kVp and mAs, using the appropriate SID
The difference is just in how the image is recorded

## Slide 13: How it works – Image Acquisition
When the PSP plate is exposed to x-rays, the x-ray photons interact with electrons in the crystals within the imaging plate
Interaction gives energy to the electrons in the crystals, trapping them in an area of the crystal called the ‘phosphor center’
  The energy transfer caused by the x-rays exposing the PSP, results in excitation of electrons into a metastable state.
  About 50% of these electrons return to their ground state immediately, emitting visible light

## Slide 14: How it works – Image Acquisition
Remaining metastable electrons are what create the latent image
  Electrons return to ground state over time
    Trapped signal can be held for hours, even for days, but the signal begins to deteriorate almost immediately – causing latent image to fade
    After 8 hours, significant signal loss may occur
    This is why it is best practice to process an exposed IR as soon as possible

## Slide 15: Europium
Is used as an activator for the phosphor
Role is to capture some of the energy during the interaction process
  Without it, there would be no latent image
When the x-rays interact with the europium activated barium fluorohalide of the phosphor layer, some light is emitted immediately, but because of the europium, some of the x-ray energy (trapped electrons) will be stored as the latent image

## Slide 16: Europium
Quantity and distribution of the trapped electrons is proportional to the x-ray exposure received
During processing, the energy of the trapped electrons is released by exposure to a laser
  Called photostimulable luminescence

## Slide 17: How it works - Stimulation
After the image has been taken, the next step is processing through the CR Reader
1st step is stimulation
  Finely focused beam of infrared light is directed at the PSP
    Used to be Helium laser beam, but now usually a solid state laser is used
    Beam diameter is 50-100 µm
      Determines the spatial resolution of the CR system
As the laser penetrates, it spreads
  The thicker the PSP is, the more the laser will spread

## Slide 18: How it Works – Reading/Detecting
The scan process produces lines of light intensity information that is detected by a photodetector
  Metastable electrons return to ground state
  Emits energy in the form of visible light
Light signal is detected and measured by the photodiodes of the photodetector and the latent image is made visible
  Some signal loss will occur due to:
    Scattering of emitted light
    Collection efficiency of photodetector
      Photodiodes (PD) are the photodetectors of choice for CR

## Slide 19: How it works - Erasing
Final stage is erasing
Stimulation cycle does not completely transition all electrons back into ground state
  Will still have some excited electrons
  Could result in a residual image on the IP
    “Ghosting” artifact
The phosphor is flooded with very intense white light from specially designed fluorescent lamps
This will remove any residual latent image

## Slide 20: Additionally…
PSPs should be used soon after they’ve been erased
Sensitive to radiation, can be fogged by background radiation
Laser light used to stimulate the PSP is monochromatic, the emitted light is polychromatic and of much lower intensity
Optical filters are needed to allow only emitted light to reach the photodetector, and block out the intense stimulating light

## Slide 21: CR Reader – Drive Mechanisms
After the CR Cassette is inserted into the CR reader, IP is removed for processing
Horizontal IP placement
IP fits into a precision drive mechanism which moves the IP slowly at a constant rate along the long axis
  Called the ‘slow scan’ mechanism
  Critical that the motor drive is absolutely constant
  Movement must be constant as even small fluctuations can result in banding artifacts

## Slide 22: CR Reader - Mechanisms
While the IP is moving in the ‘slow scan’, a mirror is used to deflect the laser beam back and forth across the IP
  This is the ‘fast scan’ mechanism
Both scans (slow and fast) work together so that the laser beam is blanked during re-trace
Error tolerance allowed is very small
  Only fractions of a pixel
  Edges of an image can appear ‘wavy’ in a CR reader that is out of tolerance

## Slide 23: CR Reader – Vertical Placement
Another method is with the cassette being placed vertically into reader
  This is what we have in lab
Once cassette is inserted, the IP is withdrawn downward and the cassette is scanned by a horizontal beam
IP barely leaves cassette so it is less likely to be damaged
Reduced occurrence of aliasing artifacts because the scan pattern is not parallel with grid lines

## Slide 24: CR Reader – Optical Components
Goal is to extract the information from the latent image as precisely as possible
  Laser
  Beam shaping optics
  Light collecting optics
  Optical filters
  Photodetector

## Slide 25: CR Reader - Laser
Source of stimulating light
Light spreads as it travels
A lens system keeps the beam diameter small when it is focused onto the mirror
  Less than 100 µm
When the mirror deflects the laser across the IP, it changes size and shape
  Additional beam shaping optics correct this and keep the size, shape and intensity constant

## Slide 26: CR Reader – Light Collection and Photodetector
Light that is emitted from the IP is then funneled into a fiber optic collection assembly before it is directed at the photodetector (or PMT, PD or CCD)
Before the light reaches the photodetector, it is filtered so that none of the stimulation light goes through and swamps the emitted light signal
  Emitted light is the signal
  Stimulating light is the noise
    Proper filtering improves the signal-to-noise ratio

## Slide 27: CR Reader – Computer Control
After the light signal leaves the photodetector, the output analog signal transmits to a computer system
  Light emitted from the phosphors is converted into an electrical signal
  The electrical signal is sampled and digitized to represent a specific brightness and location within the image matrix
Signal is shaped before the final image is formed and displayed
Image can be stored temporarily before it is transferred to a workstation or archived

## Slide 28: Imaging Characteristics
Process of imaging with film vs CR is not very different
Same process of positioning patient, using appropriate exposure techniques, patient care and radiation protection
The only difference is in what happens after the exposure is taken
  CR – processed through CR reader
  Film – processed through film processor

## Slide 29: Limitations of Film (Analog)
Limited dynamic range (the range of exposure intensities that an image receptor can accurately detect)
Cost
Time
Film processors – equipment and maintenance costs
  Frequent quality control procedures
Processed film is permanent, no adjustments can be made
Cannot be electronically stored or duplicated
Additional expense of storing radiographs and retrieving them

## Slide 30: Readings
Radiologic Science for Technologists 11th Ed
  Pg. 283 – 292 (Stop at Image Receptor Response Function)
Digital Radiography and PACS
  Pg. 58 - 69