Activity: Test 6 AEC Field Balance/Uniformity Procedure

## Purpose

This activity will help you identify the procedure steps for this test.

## Instructions

- Review the images below of the equipment and materials used at Edenbrook for this test.

- Review the procedure steps from the Edenbrook MRT lab.

- Compare the textbook readings and your practicum site QA manual for similarities and differences.

## Equipment

- Dosimeter

- Two Plexiglas blocks

- Calculator

## Edenbrook MRT Procedure Steps

- Place the Plexiglass blocks (stacked) over the AEC cells of the detector ensuring the AEC cells are completely covered.

- Using a 110 cm SID, centre the dosimeter and collimate to include the Plexiglass blocks entirely.

- Place the dosimeter probe on top of the blocks to one side of the centre cell, ensuring it is not obstructing any of the AEC cells.

- Set exposure factors of 100 mA and 70 kVp.

- Take one exposure for each AEC cell and combination of cells as follows:

  - Centre

  - Left

  - Right

  - Centre + left

  - Centre + right

  - Right + left

  - All

- Record the mAs, time in ms, and exposure in mR.

- Repeat step 4 with 50, 100, 125 and 160 mA and record the results.