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Activity: Test 17 Maximum Fluoroscopic Exposure Rate Calculations
Purpose
[bookmark: _GoBack]This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.
Instructions
1. Review the data for a completed Maximum Fluoroscopic Exposure Rate Test.
2. Compare the results to the standard in SC 35.
Evaluation of Data
	Mode
	kVp
	mA
	R/min
	mGy/min

	Continuous
	110
	3
	1.65
	3.1

	Pulse
	110
	0.72
	9.926
	3.8

	Low Dose
	110
	1.5
	0.8863
	4.6




Conclusion: Pass ______________      Fail ________________
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