# Activity: Test 17 Maximum Fluoroscopic Exposure Rate Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the data for a completed Maximum Fluoroscopic Exposure Rate Test.

- Compare the results to the standard in SC 35.

## Evaluation of Data

| Mode | kVp | mA | R/min | mGy/min |
| --- | --- | --- | --- | --- |
| Continuous | 110 | 3 | 1.65 | 3.1 |
| Pulse | 110 | 0.72 | 9.926 | 3.8 |
| Low Dose | 110 | 1.5 | 0.8863 | 4.6 |

Conclusion: Pass ______________      Fail ________________