## Slide 1: APPH 226: Grids
Objectives: 5.3, 5.4, 9.3
Notes: Bring grids to class

## Slide 2: Grid History
Invented in 1913 by Gustave Bucky
Is the most effective means for limiting scatter radiation that reaches the IR
Improvements have been made to it, but the basic principle has not changed
Consists of a radiopaque material alternating with radiolucent material

## Slide 3: Grid
Approximately ¼ inch thick and ranges from 8x10 inches to 17x17 inches
Has very thin lead strips inside with radiolucent interspaces
Purpose is to absorb scatter from the patient
Goes in between the patient and IR to improve image contrast

## Slide 4: Grids
As scatter leaves the patient, much of it is directed towards the IR
  Not great for image quality – image fog
Ideally a grid would absorb (clean up) all scatter photons and only allow transmitted photons to pass to the IR
This does not happen, but with proper use, the image contrast can be greatly improved

## Slide 5: Grid Construction - Strips
Grid strips (lines) have a precise height, thickness and space between them
  Typically made of lead because:
    Inexpensive
    Easy to shape
    High atomic number and mass density

## Slide 6: Grid Construction – Interspace
Purpose: to maintain precise separation between the lead strips
Aluminum is commonly used for interspace material
  Nonhygroscopic - does not absorb/retain moisture
  Easier to manufacture
Plastic fiber also used
  However, there is less absorption of the beam at low kVps
Lead lines and interspace material are then covered with an aluminum front and back panel

## Slide 7: Grid Construction
Grids are described by grid frequency and grid ratio
Grid Frequency
  Number of lead lines per unit length (inches/cm/both)
  Commonly in the 25-45 line/cm range
Grid Ratio
  Ratio of the height of the lead strips to the distance between them

## Slide 8: Grid Frequency
When using a stationary high frequency grid, the resultant images will show less distinct grid lines
Grids with higher frequency have thinner lead strips

## Slide 9: Grid Ratio
Defined as:
Ratio of the height of the lead strips to the distance between them
Grid Ratio = h/d
h= lead strip height D= interspace width
Grid ratio plays a major part in the ability of a grid to improve image contrast

## Slide 10: Grid Ratio
Ratios range from 4:1 to 16:1
  In general x-ray, typically 8:1 or 10:1 used
High ratio grids remove/clean up more scatter than lower ratio grids
  Increase radiographic contrast

## Slide 11: Grid Ratio
The higher the grid ratio, the straighter the scatter must be in order to pass through to the IR
Scatter must have alignment very close to primary to be able to pass through interspaces
Meaning higher ratio grids:
  Remove more scatter
  Require more accurate positioning of tube
  Prone to increased grid errors

## Slide 12: Grid Cut Off
What is it?
  Absorption of the primary beam (transmitted x-rays) by the lead strips
    ‘cut off’ from reaching IR
  Undesirable
  Loss of radiographic density
  Can be partial/complete
  Can occur with any type of grid
  Causes ‘grid errors’

## Slide 13: Grid Patterns
Refers to the pattern of the lead lines of a grid
Two types of patterns
Linear
  Grid lines run in one direction
  Most popular
Crossed or Cross hatched
  Grid lines run in two directions

## Slide 14: Crossed or Cross Hatched Pattern
Lead lines at right angles to each other
Removes more scatter than linear because there are more lead strips oriented in two directions
Limited application as the tube cannot be angled
Very precise positioning required

## Slide 15: Crossed Grids
Disadvantages
Positioning of grid is critical and CR is critical, must be perfectly centered
Exposure technique is substantial, higher patient dose
Advantages
Much more efficient at cleaning up scatter

## Slide 16: Linear Pattern
Most commonly used because it allows for angulation along the length of the lines
Can be:
  Parallel
  Focused

## Slide 17: Parallel Grid
Simplest type of grid
Has lead strips that run parallel to one another
Easiest to manufacture
Main disadvantage: cannot avoid grid cut off on image

## Slide 18: Parallel Grid
More attenuation occurs as the x-rays approach the edge of the image receptor (IR)
Would notice greater radiographic density along center line of IR, but decreasing density towards the sides
Greater grid cut off will occur with shorter SID

## Slide 19: Focused Grid
Designed to minimize grid cut off
Lead strips are angled to approximately match the angles of the divergent x-ray beam
Allows more transmitted photons to reach IR

## Slide 20: Focused Grid
Transmitted photons are more likely to pass through a focused grid to reach the IR than they are to pass through a parallel grid
Less grid cut off
Uniform density throughout image when positioned correctly

## Slide 21: Focused Grids
Clearly marked with its intended focal distance and the side that should face the x-ray tube
If used outside of the intended focal range, grid cut off will occur

## Slide 22: Focal Range
Focal Range
  Recommended range of SIDs that can be used with a focused grid

## Slide 23: Moving Grids
With stationary grids, grid lines will be produced on the images
Grid lines
  Primary x-rays absorbed within the lead strips
  Although the grid strips are very small, can still see them on an image
  Can be shown by taking an x-ray of a grid
  Usually, a high frequency grid has less obvious lines as compared with low frequency grids
    Not always, the visibility of grid lines is directly related to the width of the grid strips

## Slide 24: Moving Grids
Grid lines appearing on the image is not ideal
  Adds additional artifact
Hollis Potter came up with the idea to move the grid during the exposure
  This would make the grid lines ‘disappear’ without adding to the radiographic technique
Called a moving grid or Potter-Bucky diaphragm (bucky)

## Slide 25: Moving Grids
Two types of moving grids
Reciprocating
  Motor driven
  Moves back and forth several times during exposure
    About 2 cm distance for movement
Oscillating
  Uses spring like mechanisms with an electromagnet
  Circular motion

## Slide 26: Moving Grids
Disadvantages:
  Requires a bulky mechanism that can fail
  Increased OID due to size of mechanism
    Increases magnification of image and image blur
  Motion can be transferred to the cassette holding device, which can result in motion artifacts
Advantages far outweigh disadvantages, so moving grids are usually the technique of choice and widely implemented

## Slide 27: Grid Problems
Moving grids
  Incorrect grid installation
  If grid movement is in the same direction as the grid lines, will be ineffective and grid lines will show on the image
Stationary grids
  Most frequent error with focused grids is improper positioning
    Off level – beam angled across the grid
    Off focus – wrong SID
    Off center – not centered to CR (central ray), lateral decentering
    Upside down

## Slide 28: Off Level Grid Error
Most common grid error
Occurs when x-ray beam is angled across the grid
  Can be either tube or grid being angled
  More commonly with x-ray tube position
Incident x-rays will be angled across the grid and grid cut off will occur across entire image
Seen as loss of density over entire IR
Occurs with both focused and parallel grids

## Slide 29: Off Focus Grid Error (wrong SID)
Using SID outside of focal range for grid
Will occur if SID is greater or less than the focal range
More grid cut off will occur with higher ratio grid
The further the grid is away from the specified focal distance, the more severe the grid cut off will be
Appears as loss of density along edges of image

## Slide 30: Off Center Grid Error
Also called lateral decentering
Central Ray is not aligned from side to side with the center of a focused grid
Because of how the lead lines in the focused grid lay, the divergence of the x-ray beam does not match the angle of these strips
Appears as overall loss of density over IR

## Slide 31: Upside Down Grid Error
Grid is used upside down
Most obvious grid error
Photons will pass easily through the center of the grid as the lines are perpendicular to the IR
As the lines moving towards the sides are more angled, more photon absorption will occur
Severe grid cut off on the edges of the image, photons are completely absorbed

## Slide 32: Grids and Exposure Factors
As grid usage decreases the scatter reaching the IR, it also decreases the radiographic density
  Remember, some of the primary beam is also being absorbed
As such, when using a grid, the technical factors must be increased to compensate
Usually mAs is adjusted
Depends on the grid ratio, we use the grid conversion formula to determine the new technique, mAs

## Slide 33: Grid (Bucky) Conversion Factor
Grid Ratio | Grid Factor
No grid | 1
5:1 | 2
8:1 | 4
12:1 | 5
16:1 | 6

## Slide 34: Grid Conversion Factor (GCF)
When converting from one grid to another grid with different ratios, use the grid conversion formula.
mAs1(original) = GCF1(original grid)
mAs2 (new) GCF2 (new grid)
Carlton & Adler 4th edition pg. 547

## Slide 35: Grid Conversion Factor
A technique of 70 kVp and 5 mAs was set for a non-grid exposure. The technologist will use a 5:1 grid for the new exposure, what should the new mAs be?
mAs1(original) = GCF1(original grid)
mAs2 (new) GCF2 (new grid)
5 (mAs1) = 1 mAs2 (new) = 10
mAs2 2

## Slide 36: Grid Performance Evaluation
Ability of a grid to improve contrast is specified by:
  Contrast Improvement Factor
    Depends on amount of scatter produced
    Given values for contrast
    K = image contrast with grid
    image contrast without grid

## Slide 37: Contrast Improvement Factor
If k is 1, then no improvement is seen
Most grids will have k of 1.5 – 2.5
This means that the radiographic contrast is increased by 1.5 - 2.5 times
The higher the k factor, the greater the contrast improvement

## Slide 38: Air Gap Technique
Method to reduce scatter without using grids
  Helps to increase image contrast
IR is moved 10-15 cm away from part
Portion of the scatter will be scattered away from IR and not contribute to the image

## Slide 39: Air Gap Technique
Limitations:
  Would have to increase technique to maintain radiographic density
  Magnification of part on image
  Increased image blur
  Not as effective with higher kVp settings as more of the scattered x-rays will travel in a forward direction

## Slide 40: Readings
Bushong: Radiologic Science for Technologists 11th Edition:
  Pg 194-205