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Activity Answer: Test 15 Visual Inspection of Fluoroscopic Equipment Procedure 

Critical Thinking Question Answer

1. [bookmark: _GoBack]Why should a fluoroscopic system visual inspection occur every 6 months, while a general X-ray visual inspection occurs annually?

· More moving parts, locks, motors for power assist etc.
· Exposure switch (especially foot pedals) can malfunction or stick, increasing dose to patient and technologist, radiologist
· Lights/meters indicate when fluoroscopy is activated, both visual and auditory, if defective can increase dose to patient and staff
· More potential for things to break down 
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