# Activity - CT Linearity

Table 1: Characteristics of CT Phantom (Seeram page 201)

| Material | Density | Linear Attenuation Coefficient (cm-1) | Approximate CT Number |
| --- | --- | --- | --- |
| Polyethylene | 0.94 | 0.185 | -85 |
| Polystyrene | 1.05 | 0.196 | -10 |
| Nylon | 1.15 | 0.222 | 100 |
| Lexan | 1.20 | 0.223 | 115 |
| Acrylic (Perspex) | 1.19 | 0.229 | 130 |
| Water | 1.00 | 0.206 | 0 |
| Teflon | 2.2 | 0.479 | 990 |

| Hounsfield Unit Linearity |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Insert | Polyethylene | Teflon | Acrylic | Lexan | Nylon | Water |
| Acceptable Limits | -70 ± 40 | +950 ± 75 | +140 ± 40 | +116 ± 40 | +100 ± 40 | 0 ± 40 |
| QC Results | -56.4 | 877.4 | 129.4 | 104.4 | 94.9 | 6.7 |
| Pass/Fail | P | P | P | P | P | P |

Activity: CT Accuracy, Uniformity and Noise

ROI 2

ROI 2

ROI 4

ROI 4

ROI 1

ROI 1

ROI 5

ROI 5

ROI 3

ROI 3

Image 3 (Source: SAIT)

Image 3 (Source: SAIT)

| CT QC Chart |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Acceptable Limits | ROI size: 400 ± 50 mm2 *All values should be between -40 and +40 HU *Max ROI-Min ROI <40 |  |  |  |  | ROI *7000 ± 2000 mm2 AV: 0±40HU SD: <10 HU |  |
| QC Results | ROI 1 | ROI 2 | ROI 3 | ROI 4 | ROI 5 | 14.6 | 6.5 |
|  | 18 | 16.8 | 17.7 | 21.9 | 15.4 |  |  |
| Pass/Fail | P | P | P | P | P | P | P |

- CT Number Accuracy = All values should be between -40 & +40

- CT Uniformity limit = Max ROI – Min ROI ˂40

- CT Noise acceptable limit = SD within a large ROI must be ˂10 HU

# Activity: QC in CT

### Purpose

Objective 4.4: Identify different types of QC tests.

Recognize the parameters each QC test can evaluate.

## Instructions:

- Review the following CT QC images.

- Identify which QC test(s) was being evaluated.

Image 1: (Source https://www.ceessentials.net)

Image 1: QC parameter evaluated:

Image 2: (Source SAIT)

Image 2: QC parameters that could be evaluated:

Image 3: (Source SAIT)		            	            Image 4: (Source https://www.ceessentials.net)

Image 3: QC Parameter(s) evaluated:	               Image 4: QC Parameter evaluated:

Image 3: (Source SAIT)				           Image 4: (Source SAIT)

Image 5 (Source: http://ijcto.org/index.php/IJCTO/rt/printerFriendly/0203.9/ijcto.0203.9html)

Image 5: QC parameters evaluated:

Left		    	 Right:

Image 6 (Source: https://www.ceessentials.net)

Image 6: QC parameter evaluated:

Image 7: (Source: SAIT)		                                       Image 8: (Source: SAIT)

Image 7: QC parameter evaluated:                      Image 8: QC parameter evaluated:

Image 9: (Source: https://www.ceessentials.net)

Image 9: QC parameter evaluated: