## Slide 1: APPH 226: Generators and the X-Ray Circuit
Objectives:1.2, 1.3, 1.5, 2.1

## Slide 2: Electrodynamics Terminology
Electrodynamics
  Study of electric charges in motion
Electricity
  When an electric potential is applied to a wire, allowing the flow of electrons
Electric Circuit
  Path of electron flow from a source, through various components and back to the source
Flow of electrical current is in the opposite direction of the flow of electrons

## Slide 3: Electrodynamics
Two types of Electric Current
  Alternating Current (AC)- changes direction in cycles as the electric potential changes (negative and positive)
  Direct Current (DC) – flows only in one direction (from positive to negative, opposite the direction of electron flow)

## Slide 4: General X-Ray Circuit
Basic X-ray Circuit can be divided into:
  Main X-Ray Circuit
    Purpose: to produce x-rays
    Supplies x-ray tube with properly modified power so that it can produce x-rays
  Filament Circuit
    Purpose: to produce appropriate electron cloud
    Supplies filament of x-ray tube with properly modified power to produce appropriate thermionic cloud

## Slide 5: Components of General X-ray Circuit
Main X-Ray Circuit
  Primary Side/Circuit
    Main power switch, circuit breakers, autotransformer, timer circuit and primary side of the step up transformer
  Secondary Side/Circuit
    Secondary side of step up transformer, mA meter, rectifier bank, x-ray tube (except filaments)
Filament Circuit
  Rheostat, step down transformer, filaments

## Slide 6: Main X-Ray Circuit
Modifies the incoming line power to produce x-rays
Achieves this with a sequence of devices
Circuit needs to boost the voltage up to the range required to produce x-rays
Must permit the technologist to adjust the amperage, voltage and length of exposure

## Slide 7: Primary Circuit
Main power switch:
  An on-off switch for the unit that is connected to the power supply of the facility
Line compensator
  Used because the incoming power isn’t a consistent 220 volts
  Usually wired to the autotransformer
  Automatically adjusts the power supply to precisely 220 volts – must stay constant
Circuit breakers
  Protect against short circuits and electric shock

## Slide 8: Primary Circuit
Autotransformer
  Operates on the principle of self induction
  Only has one coil of wire around a central magnetic coil
  This coil is used as the primary and secondary coil
  Outside wires are attached at different points along the coil, and the induced voltage is adjustable depending on where the connections are made.

## Slide 9: Autotransformer
Only one core
Power brought in on one side -“primary connections”
Depending on the kVp set, different combinations of the secondary connections can increase or decrease the voltage

## Slide 10: Autotransformer
Adjustable transformer controlled by kVp selector on control panel
When a tech selects the kVp setting, they are controlling this transformer
  Determines # of turns on the secondary side to be included – which controls the output voltage
Sometimes called kVp selector
Primary purpose is to provide a voltage that will be increased by the step up transformer to produce the kilovoltage selected at the operating console

## Slide 11: Primary Circuit
Step up Transformer
  Dividing line between primary and secondary circuits
  Increases voltage from the autotransformer to the kV required for x-ray production
    Not adjustable , increases kV by a fixed amount
  Primary coil is in primary circuit, secondary coil in secondary circuit

## Slide 12: Primary Circuit
Timer circuit (exposure timer)
  Located in the primary circuit because it’s easier to control a low voltage vs a very high one (secondary circuit)
  Purpose of the device is to ‘make or break’ the voltage across the tube
  Four types:
    Synchronous timer
    Electronic timers
    mAs timers
    AEC

## Slide 13: Secondary Circuit
Begins with other side of step up transformer
An mA meter is there to monitor x-ray tube current
AC is required for the transformers to work properly, but the x-ray tube requires DC
  In the tube, electrons must always flow cathode to anode

## Slide 14: Rectifiers
Purpose: to convert AC (alternating current) to DC (direct current)
Solid state rectifier commonly used
Consists of two semiconducting crystals joined together to form a solid state diode
  P-type crystal
    ‘electron traps/holes’
    Traps are mobile
  N-type crystal
    Abundance of freely moving electrons

## Slide 15: Rectifiers
When a positive charge is placed on the p-type crystal, and negative charge on the n-type crystal, the diode will conduct electricity
Both the traps and the electrons will move across the PN junction allowing current to flow

## Slide 16: Rectifiers
When AC cycle reverses, the diode will not conduct
The traps and electrons do not move to the junction and no current is conducted

## Slide 17: Rectifiers
Solid state p-n junction conducts electricity in only one direction
  Junction is called diode
Solid state diodes are rectifiers because they conduct electric current in one direction
  One way valve
  Necessary to route the electricity through the x-ray tube correctly
Arrowhead in the diode symbol indicates direction of conventional electric current, which is opposite to the flow of electrons

## Slide 18: Why do we need rectifiers?
When the cathode is negative and the anode is positive
  Electrons will move across and x-rays will be produced
If the anode of the x-ray tube is negative and the cathode positive
  X-rays will not be produced
  Anode is not constructed to emit electrons, so there will be no electrons moving to the cathode

## Slide 19: Unrectified Voltage
Current that passes through the x-ray tube exists only during the positive half of the cycle, when the cathode is negative and anode is positive
During the negative half of the cycle, current flows from anode to cathode, but this is not possible as the anode cannot emit electrons

## Slide 20: Half Wave Rectification
Inverse voltage is removed from the supply to the x-ray tube by rectification
Done through the arrangement of rectifiers to convert AC to DC
  Need DC for x-ray tube to operate
During positive portion of AC waveform, the rectifier allows electric current to pass through tube
During negative portion of AC waveform, the rectifier does not conduct, no electric current is allowed

## Slide 21: Half Wave Rectification
Result is positive pulses with gaps when current is not being conducted
Resultant electric current is rectified because the electrons flow in only one direction
Called half wave rectification because only one half of the AC waveform appears in the output

## Slide 22: Full Wave Rectification
Major shortcoming of half wave rectification is that it wastes half of the power supply, and thus requires double the exposure time
When 4 rectifiers are used, able to utilize the negative half cycle
Inverts the negative half so that the anode is always positive

## Slide 23: Full Wave Rectification
Energy fluctuates from maximum voltage to zero – 100% ripple
This is a problem for x-ray production
X-rays produced will vary from zero to kVp selected on console

## Slide 24: Rectification
Not very effective when the energy level drops to zero
One solution is to use 3 AC waveforms at the same time, but slightly out of phase or synchronized
Each is generated independently and then phased so they are 120 degrees out of step with the next (out of step by 1/3)

## Slide 25: Three phase power
After rectifying, called three phase power
Depending on the number of rectifiers and engineering, can have ripple anywhere from 4-14%

## Slide 26: Three Phase Power
Major improvement over the single phase generator where 100% ripple meant voltage supply to the tube would produce x-rays between 0-100% of the kVp set
3 phase 6 pulse
  14% ripple
  Voltage supply to x-ray tube never falls below 86%
3 phase 12 pulse
  4% ripple
  Voltage supply to x-ray tube never falls below 96%

## Slide 27: High Frequency Generators
The next step in improving the waveform is the use of HF generators
Takes the regular 60 Hz and converts it to higher frequency and transfers it to high voltage
  Using capacitors, rectifiers, transformers
Major advantages:
  Smaller size
  Less costly to produce
  Produces a nearly constant voltage
    Minimal ripple – less than 1%

## Slide 28: Efficiency
Advantages of having less ripple are:
  Greater radiation quantity and quality produced from the more constant voltage supplied to the tube
  Quantity of photons is higher because the efficiency of x-ray production is higher when the tube voltage is high
    Reproducibility of exposure
  Fewer low energy photons

## Slide 29: In the secondary circuit
When the exposure switch is pressed, let’s summarize what happens in the primary circuit:
  Voltage goes through step up transformer, to increase to the kilovoltage selected on control panel
  kV goes through rectifiers so that it travels through the tube correctly to produce x-rays
    Large positive charge on anode
    Large negative charge on cathode
To create x-rays, need one more thing:
electrons

## Slide 30: Filament Circuit
Goal: to boil electrons out of the filament wire
The filament is a very small wire
  Must control the current that goes through it or else it could be damaged or destroyed
  Filament temperature is controlled by the filament current – measured in amperes
  A step-down transformer is used to reduce the voltage applied to the filament
mA parameter is called tube current because it reflects the rate of flow of electrons through the tube during an exposure

## Slide 31: Filament Circuit - Rheostat
This circuit begins with the rheostat
  Controlled by the mA selector on the console
  As adjustments are made to the mA on the console, the tech is actually making adjustments to the rheostat which adjusts the amount of resistance in the filament circuit, and ultimately the amount of current applied to the filament.
  Rheostat – variable resistor
    Controls filament temperature and the rate at which electrons are boiled off the filament
  Adjustments in mA = adjusting the resistance of the current in the filament circuit
    The higher the mA, the lower the resistance

## Slide 32: Filament Circuit
To provide voltage for the filament circuit, there are connections on the autotransformer
  From the autotransformer, the voltage goes to the rheostat, and then a step down transformer
Resistors are used to reduce the voltage to correspond with the selected mA
  Resistors are in fixed stations to provide specific mA ‘stations’ – example…100, 200, 300 mA etc…

## Slide 33: Filament Circuit
Rheostat works together with the exposure timer
Rheostat controls filament temperature and the electrons being emitted, and the timer determines the duration of this process
Together they determine the quantity of electrons boiled off and available to produce x-rays
Step down transformer is used to reduce the voltage applied to the filament

## Slide 34: Filament Circuit Summary
When exposure button is pressed, voltage and current is sent through the filament circuit
Quantity is controlled by rheostat to reflect the mA selected
Step down transformer is used to reduce the voltage and increase current (inverse relationship)
Current heats the selected filament and begins boiling off electrons
Time selected controls how long the voltage and current is applied to the tube

## Slide 35: Summary of Circuit Operation
On the console, select the kVp, mA, exposure time and focal spot size
Follow voltage through the primary and secondary circuit
Follow current through the filament circuit

## Slide 36: Path of Circuit
1. kVp selected adjusts the autotransformer
  Determines the number of turns on the secondary side necessary to produce a voltage to send to the step up transformer
  Step up transformer increase this voltage by a fixed amount and produces the kV selected on the console
  This transformer represents the transition from the primary (low voltage) to the secondary (high voltage) circuit

## Slide 37: Path of Circuit
2. kV is created, now it must be rectified for the x-ray tube
  Electrons must flow from cathode to anode
  Rectifiers (one way valves) will route the electricity through the tube correctly
  After passing through the rectifiers, the electricity will create a large positive charge on the anode, and a large negative charge on the cathode (focusing cup)

## Slide 38: Path of Circuit (Filament)
Filament circuit draws electricity from the autotransformer to the rheostat
Rheostat is controlled by mA selector on the console and is connected to the focal spots
Focal spot selector represents the filaments in the tube
mA station selection will also determine which filament is selected

## Slide 39: Path of Circuit (Filament)
Selected mA sets the resistance in the filament circuit
From the rheostat, electricity travels to the step down transformer
Adjusted current travels to the filament
Filament is heated up and thermionic emission occurs

## Slide 40: Path of Circuit
Now the paths join
Electron cloud created by the filament is ready and waiting
kV applied to the x-ray tube creates a large positive charge on the anode and a large negative charge on the cathode (focusing cup)
Positive charge on anode draws the electrons over, giving them tremendous kinetic energy in the process

## Slide 41: Path of Circuit
Opposites attract and in the short 1-3 cm gap between cathode and anode the electrons reach speeds of up to ½ the speed of light
Negative charge on cathode serves to keep electrons crowded together
Electrons travel to the anode and interact there to produce x-rays until the timer circuit terminates the process

## Slide 42: Falling Load Generator
Uses the shortest exposure time possible
Method: initial tube loading is higher, and drops during exposure
Results in achieving the set mAs, in a shorter exposure time
Principal use in areas where the shorter the exposure time, the better
  Interventional radiology

## Slide 43: Emission Spectrum and Generator
X-ray production varies at different part of the voltage cycle
Quantity of x-rays is greater at peak voltages than at lower voltages
Results in considerably higher x-ray emission for 3 phase or HF waveforms

## Slide 44: Emission Spectrum and Generators
As the efficiency of the voltage wave form increases, the efficiency of x-ray production increases also
High frequency generators have the lowest ripple, so they will be at peak voltage for most of the duration of the exposure
  Meaning more efficient production of x-rays (quantity) and higher energy x-rays produced (higher average)
Single phase generators have the most ripple
  Less efficient production of x-rays, lower average energy

## Slide 45: Readings
Bushong: Radiologic Science for Technologists 11th Edition
  Pg 86-102
  If you would like to review Electricity, Magnetism and Electromagnetism : Read over Chapter 4 pg 61-81