Activity Answers: Test 2 kVp Accuracy Test

## Evaluation of Data

| Set kVp | Set mAs | Focal Spot | Actual kVp | kVp diff (%) | Pass | Fail |
| --- | --- | --- | --- | --- | --- | --- |
| 50 | 10 | Small | 49.43 | -1.14 | X |  |
| 50 | 10 | Large | 49.62 | -0.76 | X |  |
| 60 | 10 | Small | 59.63 | -0.616 | X |  |
| 60 | 10 | Large | 59.45 | -0.916 | X |  |
| 70 | 10 | Small | 69.69 | -0.442 | X |  |
| 70 | 10 | Large | 69.65 | -0.5 | X |  |
| 80 | 10 | Small | 80.66 | 0.8251 | X |  |
| 80 | 10 | Large | 80.48 | 0.61 | X |  |
| 100 | 10 | Small | 101.22 | 1.22.5 | X |  |
| 100 | 10 | Large | 101.44 | 1.443 | X |  |

- Determine the percent (%) difference between set and measured kVp values.

% Difference = (Actual kVp) – (Set kVp) x 100

(Set kVp)

- Record the % difference for each set kVp and focal spot size.

- Compare the results of the % differences to the standard in SC 35.

- The measured kVp shall have a value within +/- ______10%_____ of the set kVp value.

Conclusion: Pass _______X______      Fail ________________

## Critical Thinking Questions

- Kilovoltage settings tend to drift over time. What could be a cause for this?

- Tubing aging

- Variations in line voltage supplying the X-ray generator

- Faulty high-voltage cable

- Problems with the autotransformer/kVp selection circuity

- Inaccurate reading of the detector

- In this test, a digital kVp meter was used. What other type of equipment can be used to perform this test?

  - Wisconsin Test Cassette

  - Ardan and Crook’s Cassette

Which is more accurate and why?

  - The digital meters are usually more accurate and easier to use because when they are exposed, the measured kVp appears automatically with a light-emitting diode (LED) readout. More on page 94 of text

- Why does this test need to be done at multiple kVp settings?

  - this test conforms the peak voltage at various kilovolt (peak) stations available for a particular generator.

  - should be done in 10-20 kVp increments starting at 50 kVp

  - variations in output can be caused by variations in the line voltage supplying the X-ray generator, by faulty high-voltage cables, or by problems with the autotransformer/kilovolts (peak) selection circuitry.  Papp, page 100.