## Slide 1: MRAD 251: CT Theory 1
Outcome 2 (Part 1):
Analyze computed tomography (CT) system design and components

## Slide 2: Objectives
Apply terminology specific to CT equipment
Characterize the various acquisition components that comprise a CT scanner
Outline the differences of the seven generations of CT scanners
Describe the applications of CT fluoroscopy, dual source CT, CT simulation and CT in nuclear medicine
Explain the components and operation of contrast medium injection devices

## Slide 3: Recap: CT Acquisition
The first step in creating a CT Image.
Describes the process of collecting attenuation data
Requires the use of specialized pieces of equipment
  Must be capable of capturing x-ray energy, and converting the energy into other useable forms
  Can be used separately, or in conjunction to complete the task of creating a diagnostic CT image

## Slide 4: Objectives 2.1 & 2.2
Basic CT Equipment:
Couch/Table
Operating Console
Workstation
Contrast media injector
Accessory Equipment
Gantry

## Slide 5: 1. Couch/Table
The patient is positioned on this piece of equipment.
Moves the patient through the radiation beam
Physical characteristics of a CT Table:
Narrow rectangular shape
Usually concave
Made of carbon fiber
  Strong & rigid material
Low absorption qualities
Electrical and mechanical controls for table movement
  Example: Quick release locks
Attachment areas for accessory equipment

## Slide 6: Table Movements
The table is capable of moving in two planes.
Vertical (Y-axis):
Enables the technologist to work ergonomically
Facilitates an easier patient dismount
Allows the patient to be positioned within the gantry isocenter
Horizontal (X-axis):
Moves the patient in and out of the gantry
  Enabling irradiation of an anatomical area of interest
When the table moves a specified distance in this plane, it is known as (aka. feed, step, or index)

## Slide 7: Table Considerations
Table motion can affect consistency of scanning and locating pathologies.
When a patient is positioned within the isocenter of the gantry, table position
should be
  A digital readout displays table position relative to the gantry (aka. Scan point)
Identifies location of irradiation
Should be manually set to zero by the technologist
Limitations:
Weight restrictions
  Exceeding weight limits could affect precision of Pitch
•
  Horizontal motion limits
  Must be considered to prevent repositioning

## Slide 8: 2. Operating Console
The pieces of equipment that allow the technologist to input factors affecting the scan prior to acquisition.
Consists of a computer, keyboard and multiple monitors
  Keyboard enables:
    Communication with the scan room
    Manipulation of gantry tilt
    Table movement
    Irradiation
  Monitors are usually defined as:
    Scan monitor, or
    Display monitor

## Slide 9: Operating Console Continued
A enables the selection of various factors prior to acquisition, such as:
Patient demographics
Scan protocol
Scan parameters
A enables:
Image viewing
Post-processing functions
The operators console stores raw data.
Storage time is limited

## Slide 10: 3. Workstation
Utilized to increase workflow efficiency within a CT department.
Enables simultaneous scan acquisition and image review of different patients
Consists of a computer, keyboard and multiple monitors
  Keyboard is only capable of inputting data into a computer
Allows the technologist to post-process images
  Can only access and manipulate image data

## Slide 11: 4. Contrast Media Injector
A mechanical device, used to safely administer large amounts of contrast media through an IV.
Used to enhance subject contrast
Can be programmed to ensure:
  Precise volumes, and injection rates are used
    Creates consistency of image appearance
Features:
  Audible and visual alerts
    Used to help identify potential problems
      Example: Pressure limits
  Syringes
    Used to store contrast media and saline separately
    Syringes containing contrast require a “warming sleeve”

## Slide 12: Objective 2.5
Contrast medium injection devices typically comprise of:
Syringes:
  One or Two
    Single syringe devices = Single head injector
    Two syringes = Dual head injector systems
      One syringe filled with contrast, the other with saline
Warming sleeves
  Used to maintain the desired temperature of contrast media
Console
  Enables selection, and monitoring, of pre-programmed protocols, and
other contrast administration factors

## Slide 13: 5. Accessory Equipment
Pieces of equipment that are only used when needed. Designed to improve patient care, and image quality. Accessory equipment commonly used in CT:
Specialized headrest
  Assists with patient positioning for head scans
Sponges
  Help maintain patient position and increase comfort
Immobilization straps
  For patient safety, compliance, and motion reduction
Shielding
  Used for patient protection

## Slide 14: 6. Gantry
Piece of equipment a patient moves through.
Basic Characteristics:
~ 70→90cm diameter aperture
  “donut” shaped
Tilts up to 30°
  Cephalad & caudad
Has positioning lights and control panel(s)
  Controls angulation
  Enables accurate patient positioning
    Isocenter
  Has a microphone for patient communication

## Slide 15: Gantry Continued
Houses many of the components needed to produce and detect x- rays for image formation:
DAS
Slip Rings
*Generator
X-ray Tube
Filtration
Collimators
Detector Array

## Slide 16: a) DAS
Aka. Data acquisition system
Measures the number of photons that strike the detector array (aka. transmitted radiation beam)
  Via sampling
    Reduces creation of image artifacts
Converts measurements into a digital signal (binary)
  Using an (analog to digital converter)
Sends the digital data to the computer for reconstruction

## Slide 17: b) Slip Rings
Electromechanical devices
A brush-like apparatus
Provides continuous electrical power & electronic communication across a rotating surface
  Enables the gantry frame to rotate continuously
  Eliminates the need to straighten twisted system cables

## Slide 18: c) Generator
Conventional units used 3-phase generators
Large
  Housed outside of the gantry
Modern units use high frequency generators
Compact
  Housed within the gantry
Better efficiency
Can be stationary or rotating
High-frequency inverter circuit
  Enables high kV settings
  Produces less ripple

## Slide 19: d) X-ray Tube
Designed to decrease heat loads:
High anode rotation speeds
Larger and thicker anodes
Metal envelopes
  Prevent arcing
Dual focal spots
  Common sizes: 0.5mm & 1mm
Power supplied by high frequency generators also decrease heat loading
  Common range: 120–140 kVp

## Slide 20: e) Filtration
Shapes the energy distribution across an x-ray beam, by absorbing low energy photons.
Produces a more homogeneous beam
  By increasing the average energy of the beam
    Reduces patient dose
Detectors operate more consistently with a uniform beam
  Increases accuracy of linear attenuation coefficient measurements
    Minimizes the appearance of artifacts caused by beam hardening

## Slide 21: Filtration Continued
Filtration can be classified in two ways.
Inherent
  Built into the system
  Can not be changed
  Applied to all exams
Added (aka. compensating)
  Located between the tube and patient
  Applied to anatomically specific exams
  Have special shapes, that match beam divergence and patient shape
    Categorized by shape (example: Bow tie)

## Slide 22: f) Collimation
Small, thin pieces of lead that resemble shutters.
Located at two different locations:
  Pre-patient (aka. source)
    Decrease patient dose
  Post-patient (aka. pre-detector)
    Maintains beam width at the detector(s)
    Decreases the amount of scatter reaching the detector(s)
Restrict the x-ray beam to anatomy of interest (aka. narrows the x- ray beam)
  Contributes to slice thickness in MDCT
  Controls slice thickness in SDCT
Controls voxel length (z-axis)
  (common slice thicknesses: 0.5mm→10mm )

## Slide 23: g) Detectors
Measure transmitted radiation photons.
Ultimately convert measurements into an electrical signal proportional to radiation intensity
Scatter detection is affected by:
  Placement of detectors
    Spacing bars contribute to distance measurements
  Shape
    Arc configurations vs. linear configurations
  Size of the detector aperture
    Short, wide detectors absorb more scatter
    ↓aperture size = ↑spatial resolution

## Slide 24: Detectors Considerations
Detector quality can be determined by four characteristics:
Quantum Detection Efficiency (QDE)
  Capture, absorption, and conversion of radiation data
High Stability
  Response to radiation must be consistent
  Determined by: Calibration requirements
Fast Response Time
  Lag time is undesirable
Wide Dynamic Range
  Ability to measure different signals

## Slide 25: Ionization (Xenon Gas) Detectors
Used in conventional scanner models.
Not used in MDCT/MSCT
Converts x-ray photon energy into electrical energy
Advantages:
Cost
Stability
Response time
Disadvantages:
Pressure
Space
  Caused by Tungsten plates
50-87% QDE
  Due to aluminum casing

## Slide 26: Solid State/Scintillation Detectors
Type of detector used in modern systems.
Used in MDCT/MSCT
“2-step” energy conversion
•
  Systems use a photocathode or photodiode to convert light into electrical
energy
Advantages:
Absorption efficiency (94-100%)
Disadvantages:
Afterglow
Temperature and moisture sensitive
Spectral Matching

## Slide 27: Solid State/Scintillation Detectors
This type of detector requires the following characteristics:
High light output
  Light energy should = captured X-ray energy
    Requires the use of an
High x-ray stopping power
  Ideally, detectors could absorb all photons that strike its surface, regardless of energy level
Good spectral match
  Compatibility with the color of light emitted
Fast response time
  Low afterglow (nanoseconds)

## Slide 28: Detector Arrays
CT systems use multiple detectors during scan acquisition.
Detector arrangements determine how many image slices can be recorded with a single gantry rotation
Single detector row systems (aka. SDCT):
Detectors are arranged side-by-side, in a single row
  They have a wide Z-axis
    Creates a fan-beam
Collimation controls slice thickness
  Largest allowable is less than the full detector width
Each 3600 rotation produces one image slice

## Slide 29: Detector Arrays Continued
Multi-detector row systems (aka. MDCT):
Many detectors are arranged within a row
  They have narrow Z-axis widths; creating a cone-beam
Enables the production of multiple slices per 3600 rotation
:
  number of slices and slice thickness
    ↓overall scan times
Two main configurations affect how detector elements can be combined to determine the number of slices produced, and their thickness, during each rotation of the tube

## Slide 30: MDCT Detector Configurations
Matrix Array
  Aka. Fixed Array, or Uniform Array
    Isotropic voxels are produced
      Goal is to improve spatial resolution
      Slice thickness is symmetrical
Adaptive Array
  Aka. Non-uniform Array, or Hybrid Array
    Anisotropic voxels are produced
      Slice thickness is variable
        Thinner in the center, wider at the periphery
Note: Image slice thickness cannot be smaller than the slice thickness selected
during acquisition.

## Slide 31: MDCT Detector Advantages
Data acquisition is faster than SDCT systems.
More slices are obtained in a single 3600 rotation
  4 → 320 slices
↑anatomical coverage in a shorter scan time
  Breath holds can be ↓
  ↓ probability of imaging motion
Ability to combine detector elements.
Can manipulate slice thickness after acquisition
  Aka. Retrospective slice thickness
Thinner slices minimize partial volume averaging

## Slide 32: Retrospective Slice Thickness
Images can be manipulated to create different slice thicknesses than those used for acquisition.
•
Image slice thickness ≥ scan slice thickness
  Slices can be combined to create larger image slices
  Example: Two 1mm slices are acquired, and are reconstructed to
display one single 2mm slice
Does not affect patient dose
  Re-uses raw data

## Slide 33: Objective 2.3
Technological advancements have emerged, in response to limitations associated with CT scanners.
Scanners are classified into generations
  Scanner generation is determined by the design of the x-ray tube and detector(s)
  Organized chronologically
    Example: 1st Generation scanners are older than 3rd Generation scanners
7th Generation – prototype
  Flat panel detector CT systems
  Potential future use for breast imaging

## Slide 34: Objective 2.4
A variety of exams can be performed using CT systems, due to the 3D nature of data acquisition, and post-processing capabilities.
Additional CT studies include:
  Angiography
  Fluoroscopy
  Screening
  Nuclear medicine
  Simulation

## Slide 35: Things to take away
Equipment characteristics
  Console vs. Workstation
  Table
    Incrementation
    Scan point
    Scannable range
  Gantry controls & components
Detectors
  Types
  Configurations
CT Generations
Other applications of CT