Activity: Test 7 Automatic Exposure Control (AEC) Density Control Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the sample data for a completed AEC Density Control Test.

- Complete the calculations below.

## Evaluation of Data

| Density Setting | mAs | mR |
| --- | --- | --- |
| -2 | 2.0 | 12.2 |
| -1 | 2.6 | 15.57 |
| 0 | 3.2 | 19.35 |
| +1 | 4.1 | 24.14 |
| +2 | 5.0 | 30.44 |

- Determine the percent difference in mR and mAs between adjacent density control settings.

(E.g., between -2 and -1 density setting.)

% Difference = (-2 setting) – (-1 setting)   x 100

(-1 setting)

mR

% = (           ) – (          )   x 100                [Density setting -2 and -1]

(         )

[Density setting -1 and 0]

[Density setting 0 and +1]

[Density setting +1 and +2]

mAs

% = (           ) – (          )   x 100                 [Density setting -2 and -1]

(         )

[Density setting -1 and 0]

[Density setting 0 and +1]

[Density setting +1 and +2]

- Compare results to SC 35 and Papp, 2015, pp. 111–113.

- SC 35 states _________________________ should be used. The AHS standard states there should be a difference of 15–25% between adjacent settings.

Conclusion: Pass ______________      Fail ________________